## InAir
In Air Simulations for Large Collimations |  | Fluka Src: CTBOWHEEL |  |  |  | GeV/cm^2 | [5 empty] | AIR |  |  | GeV/g
With Bowtie effect and Heel effect. No Overbeaming Effect | [5 empty] | Energy FLUENCE | [5 empty] | (Linear Interpolation) |  |  | Dose
| ENERGY Bounds (GeV) | Mid-point Energy (GeV) | Mid-point Energy (KeV) | CTDI100 ¦16cm Col | QA ¦16cm Col | CTDI100 ¦14cm Col | QA ¦14cm Col | CTDI100 ¦12cm Col | QA ¦12cm Col | CTDI100¦80cm Col | QA ¦80cm Col | Mass Energy attenuation cm^2/g | CTDI100 ¦16cm Col | QA ¦16cm Col | CTDI100 ¦14cm Col | QA ¦14cm Col | CTDI100 ¦12cm Col | QA ¦12cm Col | CTDI100¦8cm Col | QA ¦8cm Col
1 | 0.0001505 | 0.000148 | 148 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.024892 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
2 | 0.0001455 | 0.000143 | 143 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.024721 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
3 | 0.0001405 | 0.000138 | 138 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.02455 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
4 | 0.0001355 | 0.000133 | 133 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.024379 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
5 | 0.0001305 | 0.000128 | 128 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.024208 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
6 | 0.0001255 | 0.000123 | 123 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.024037 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
7 | 0.0001205 | 0.000118 | 118 | 8.2027e-10 | 8.1692e-10 | 8.4295e-10 | 8.549e-10 | 9.6663e-10 | 1.1438e-09 | 1.2217e-09 | 1.3829e-09 | 0.023866 | 1.9576e-11 | 1.9496e-11 | 2.0118e-11 | 2.0403e-11 | 2.3069e-11 | 2.7297e-11 | 2.9158e-11 | 3.3004e-11
8 | 0.0001155 | 0.000113 | 113 | 2.4272e-09 | 2.3726e-09 | 2.688e-09 | 2.8075e-09 | 3.159e-09 | 3.1042e-09 | 3.906e-09 | 4.6933e-09 | 0.023695 | 5.7513e-11 | 5.6218e-11 | 6.3691e-11 | 6.6524e-11 | 7.4852e-11 | 7.3552e-11 | 9.2551e-11 | 1.1121e-10
9 | 0.0001105 | 0.000108 | 108 | 3.9119e-09 | 4.0462e-09 | 4.4739e-09 | 4.5504e-09 | 4.9753e-09 | 4.8864e-09 | 6.3581e-09 | 7.1256e-09 | 0.023524 | 9.2021e-11 | 9.5181e-11 | 1.0524e-10 | 1.0704e-10 | 1.1704e-10 | 1.1495e-10 | 1.4956e-10 | 1.6762e-10
10 | 0.0001055 | 0.000103 | 103 | 5.2842e-09 | 4.8499e-09 | 5.9512e-09 | 6.1437e-09 | 6.948e-09 | 6.7286e-09 | 8.8485e-09 | 1.0116e-08 | 0.023353 | 1.234e-10 | 1.1326e-10 | 1.3898e-10 | 1.4347e-10 | 1.6225e-10 | 1.5713e-10 | 2.0664e-10 | 2.3624e-10
11 | 0.0001005 | 9.8e-05 | 98 | 6.5484e-09 | 6.6899e-09 | 7.3165e-09 | 7.304e-09 | 8.5091e-09 | 8.566e-09 | 1.0835e-08 | 1.1948e-08 | 0.023332 | 1.5279e-10 | 1.5609e-10 | 1.7071e-10 | 1.7042e-10 | 1.9853e-10 | 1.9986e-10 | 2.5279e-10 | 2.7877e-10
12 | 9.55e-05 | 9.3e-05 | 93 | 7.6842e-09 | 7.6067e-09 | 8.648e-09 | 8.6024e-09 | 1.0082e-08 | 9.6738e-09 | 1.2609e-08 | 1.426e-08 | 0.023537 | 1.8086e-10 | 1.7904e-10 | 2.0355e-10 | 2.0247e-10 | 2.3731e-10 | 2.2769e-10 | 2.9678e-10 | 3.3564e-10
13 | 9.05e-05 | 8.8e-05 | 88 | 8.4823e-09 | 8.5597e-09 | 9.7623e-09 | 9.5151e-09 | 1.1208e-08 | 1.0877e-08 | 1.4303e-08 | 1.6626e-08 | 0.023742 | 2.0139e-10 | 2.0322e-10 | 2.3178e-10 | 2.2591e-10 | 2.6611e-10 | 2.5823e-10 | 3.3959e-10 | 3.9473e-10
14 | 8.55e-05 | 8.3e-05 | 83 | 9.5987e-09 | 9.4702e-09 | 1.0747e-08 | 1.1021e-08 | 1.2168e-08 | 1.2219e-08 | 1.5438e-08 | 1.7828e-08 | 0.023947 | 2.2986e-10 | 2.2678e-10 | 2.5737e-10 | 2.6391e-10 | 2.9139e-10 | 2.9261e-10 | 3.6969e-10 | 4.2692e-10
15 | 8.05e-05 | 7.8e-05 | 78 | 1.0269e-08 | 1.0237e-08 | 1.1419e-08 | 1.0822e-08 | 1.3496e-08 | 1.2892e-08 | 1.6672e-08 | 1.909e-08 | 0.024704 | 2.5369e-10 | 2.5288e-10 | 2.821e-10 | 2.6735e-10 | 3.3341e-10 | 3.1849e-10 | 4.1187e-10 | 4.7159e-10
16 | 7.55e-05 | 7.3e-05 | 73 | 1.0481e-08 | 1.042e-08 | 1.1876e-08 | 1.1129e-08 | 1.3471e-08 | 1.3458e-08 | 1.7417e-08 | 1.9423e-08 | 0.026289 | 2.7553e-10 | 2.7392e-10 | 3.1221e-10 | 2.9257e-10 | 3.5413e-10 | 3.5379e-10 | 4.5786e-10 | 5.1061e-10
17 | 7.05e-05 | 6.8e-05 | 68 | 1.5253e-08 | 1.4694e-08 | 1.7499e-08 | 1.6892e-08 | 2.0083e-08 | 1.9654e-08 | 2.5258e-08 | 2.9093e-08 | 0.027874 | 4.2516e-10 | 4.0959e-10 | 4.8776e-10 | 4.7084e-10 | 5.5979e-10 | 5.4784e-10 | 7.0404e-10 | 8.1095e-10
18 | 6.55e-05 | 6.3e-05 | 63 | 1.3552e-08 | 1.2779e-08 | 1.5471e-08 | 1.5161e-08 | 1.7832e-08 | 1.7693e-08 | 2.2594e-08 | 2.5339e-08 | 0.029459 | 3.9924e-10 | 3.7646e-10 | 4.5577e-10 | 4.4662e-10 | 5.2531e-10 | 5.2121e-10 | 6.6561e-10 | 7.4646e-10 |  | CTDI100 ¦16cm Col | QA ¦16cm Col | CTDI100 ¦14cm Col | QA ¦14cm Col | CTDI100 ¦12cm Col | QA ¦12cm Col | CTDI100¦8cm Col | QA ¦8cm Col
19 | 6.05e-05 | 5.8e-05 | 58 | 2.0414e-08 | 1.9769e-08 | 2.2995e-08 | 2.2941e-08 | 2.6369e-08 | 2.6592e-08 | 3.4012e-08 | 3.7904e-08 | 0.032524 | 6.6394e-10 | 6.4296e-10 | 7.479e-10 | 7.4613e-10 | 8.5763e-10 | 8.6488e-10 | 1.1062e-09 | 1.2328e-09 | Measurement | 5573 | 6024 | 5557 | 6013 | 5538 | 5995 | 4720 | 5935
20 | 5.55e-05 | 5.3e-05 | 53 | 1.2702e-08 | 1.2507e-08 | 1.4583e-08 | 1.422e-08 | 1.6578e-08 | 1.5904e-08 | 2.1175e-08 | 2.3673e-08 | 0.037809 | 4.8025e-10 | 4.7287e-10 | 5.5136e-10 | 5.3763e-10 | 6.2679e-10 | 6.013e-10 | 8.0062e-10 | 8.9504e-10 | Fluka | 6010.1887 | 5878.0764 | 5963.4634 | 5830.8083 | 6036.987 | 5906.2936 | 5228.0552 | 5903.3613
21 | 5.05e-05 | 4.8e-05 | 48 | 1.1366e-08 | 1.1285e-08 | 1.2857e-08 | 1.2142e-08 | 1.4751e-08 | 1.47e-08 | 1.8581e-08 | 2.1483e-08 | 0.04645 | 5.2793e-10 | 5.242e-10 | 5.972e-10 | 5.6397e-10 | 6.8516e-10 | 6.8279e-10 | 8.631e-10 | 9.9787e-10
22 | 4.55e-05 | 4.3e-05 | 43 | 9.1858e-09 | 8.8341e-09 | 1.0196e-08 | 1.0252e-08 | 1.1875e-08 | 1.1652e-08 | 1.5065e-08 | 1.7408e-08 | 0.060125 | 5.523e-10 | 5.3115e-10 | 6.1304e-10 | 6.1638e-10 | 7.1397e-10 | 7.0059e-10 | 9.0579e-10 | 1.0467e-09
23 | 4.05e-05 | 3.8e-05 | 38 | 6.3543e-09 | 6.2995e-09 | 7.1795e-09 | 6.8506e-09 | 8.2575e-09 | 8.044e-09 | 1.0509e-08 | 1.1442e-08 | 0.085404 | 5.4269e-10 | 5.38e-10 | 6.1316e-10 | 5.8507e-10 | 7.0522e-10 | 6.8699e-10 | 8.9752e-10 | 9.7723e-10
24 | 3.55e-05 | 3.3e-05 | 33 | 3.3034e-09 | 3.21e-09 | 3.7125e-09 | 3.5873e-09 | 4.3111e-09 | 3.9843e-09 | 5.391e-09 | 5.9261e-09 | 0.12809 | 4.2313e-10 | 4.1116e-10 | 4.7553e-10 | 4.595e-10 | 5.522e-10 | 5.1035e-10 | 6.9053e-10 | 7.5906e-10
25 | 3.05e-05 | 2.8e-05 | 28 | 9.8789e-10 | 9.3278e-10 | 1.1283e-09 | 1.111e-09 | 1.3382e-09 | 1.2438e-09 | 1.666e-09 | 1.7854e-09 | 0.23074 | 2.2795e-10 | 2.1523e-10 | 2.6035e-10 | 2.5636e-10 | 3.0878e-10 | 2.87e-10 | 3.8441e-10 | 4.1196e-10
26 | 2.55e-05 | 2.3e-05 | 23 | 1.2062e-10 | 1.1734e-10 | 1.2189e-10 | 1.1008e-10 | 1.4778e-10 | 1.2655e-10 | 1.6027e-10 | 2.0006e-10 | 0.42334 | 5.1063e-11 | 4.9675e-11 | 5.1601e-11 | 4.6601e-11 | 6.2561e-11 | 5.3574e-11 | 6.7849e-11 | 8.4693e-11
27 | 2.05e-05 | 1.8e-05 | 18 | 2.85e-12 | 8.3e-13 | 2.1e-12 | 4e-14 | 3.41e-12 | 1.93e-12 | 3.64e-12 | 1.14e-12 | 0.85694 | 2.4423e-12 | 7.1126e-13 | 1.7996e-12 | 3.4278e-14 | 2.9222e-12 | 1.6539e-12 | 3.1193e-12 | 9.7691e-13
28 | 1.55e-05 | 1.3e-05 | 13 | 4.9e-13 | 5e-13 | 7.4e-13 | 5.2e-13 | 5.5e-13 | 1.31e-12 | 1.08e-12 | 1.88e-12 | 2.6972 | 1.3216e-12 | 1.3486e-12 | 1.9959e-12 | 1.4025e-12 | 1.4835e-12 | 3.5333e-12 | 2.913e-12 | 5.0707e-12
29 | 1.05e-05 | 8e-06 | 8 | 5.2e-13 | 8.3e-13 | 5.8e-13 | 5.1e-13 | 3e-13 | 4.9e-13 | 7.9e-13 | 1.19e-12 | 9.446 | 4.9119e-12 | 7.8402e-12 | 5.4787e-12 | 4.8175e-12 | 2.8338e-12 | 4.6285e-12 | 7.4623e-12 | 1.1241e-11
30 | 5.5e-06 | 3e-06 | 3 | 1.7e-13 | 1.8e-13 | 1.4e-13 | 1.7e-13 | 1.3e-13 | 1.7e-13 | 2.3e-13 | 3.4e-13 | 161.4 | 2.7438e-11 | 2.9052e-11 | 2.2596e-11 | 2.7438e-11 | 2.0982e-11 | 2.7438e-11 | 3.7122e-11 | 5.4876e-11
| 5e-07 |  |  |  |  | 1.7947e-07 | 1.7602e-07 | 2.0653e-07 | 2.0315e-07 | 2.6203e-07 | 2.9675e-07 | 0.36048 | 5.9164e-09 | 5.7863e-09 | 6.6713e-09 | 6.5229e-09 | 7.6837e-09 | 7.5174e-09 | 9.7428e-09 | 1.1001e-08
[12 empty] | Dose (mGy/particle) | 9.4791e-13 | 9.2707e-13 | 1.0689e-12 | 1.0451e-12 | 1.2311e-12 | 1.2044e-12 | 1.561e-12 | 1.7626e-12
| Constants | [10 empty] | Exposure mR /particle | 1.0821e-10 | 1.0583e-10 | 1.2202e-10 | 1.193e-10 | 1.4053e-10 | 1.3749e-10 | 1.7819e-10 | 2.0121e-10
[7 empty] | 3.22e-07 | 2e-07
1.61
[12 empty] | Fluka Exp @ 300 mAs | 6010.1887 | 5878.0764 | 5963.4634 | 5830.8083 | 6036.987 | 5906.2936 | 5228.0552 | 5903.3613
[11 empty] | (Measured at 1s Rotation) | Measured Exp @ 300 mAs | 5573 | 6024 | 5557 | 6013 | 5538 | 5995 | 4720 | 5935
[12 empty] | Rel. Error | 0.078448 | 0.024224 | 0.073144 | 0.0303 | 0.090102 | 0.014797 | 0.10764 | 0.0053309
Source Strength @ 300 mAs | [11 empty] | Correction Factor | 0.92726 | 1.0248 | 0.93184 | 1.0312 | 0.91735 | 1.015 | 0.90282 | 1.0054
Gev/G to mGy | 0.00016022 |  |  |  | Expected Src Strengths
Collimation (cm) | Bowtie | Bowtie+FilmHEEL | Bowtie+ChamberHEEL | Bowtie+newFilm | CTDI | QA | Mean |  |  |  |  | Old-Film | Day 2
16 | 5.6676e+13 | 5.5542e+13 | 5.3411e+13 | 5.0733e+13 | 5.1502e+13 | 5.6921e+13 |  |  |  |  | Collimation | Fluka 1.9cm chamber | Measured 1.9cm Chamber
14 | 5.021e+13 | 48874646514455 |  |  | 4.5543e+13 | 1.0312 |  |  |  |  | 16 | 5878.0764 | 6024
12 | 43345505620919 | 4.2958e+13 |  |  | 0.91735 | 4.3603e+13 |  |  |  |  | 14 | 5830.8083 | 6013
8 | 2.9669e+13 | 2.9339e+13 |  |  | 2.6488e+13 | 2.9497e+13 |  |  |  |  | 12 | 5906.2936 | 5995 | 1.0784 |  | Average Correction Factor (CTDI Chamber)
[11 empty] | 8 | 5903.3613 | 5935 |  |  | 0.91982
160 Collimation
|  |  | CTDI160 NEW Film | CTDI160 OlD Film | QA160 New Film | QA160 Old Film | QA chamber
| Day2 Correction Factor for source strength |  | 0 | 0 | 0 | 0 | 0
| 1.008 |  | 0 | 0 | 0 | 0 | 0 |  |  |  |  | 160 Collimation - QA chamber Measurement
|  |  | 0 | 0 | 0 | 0 | 0 |  |  |  |  | Day 1 | 5978
|  |  | 0 | 0 | 0 | 0 | 0 |  |  |  |  | Day 2 | 6024
|  |  | 0 | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 0 | 0 | 0 |  |  |  |  | 160 Collimation - Heel differences - Day 1 Source Strength
|  |  | 8.2089e-10 | 8.2027e-10 | 8.985e-10 | 8.1692e-10 | 7.1086e-10 |  |  |  |  | old Film | 5878.0764
|  |  | 2.5839e-09 | 2.4272e-09 | 2.538e-09 | 2.3726e-09 | 2.5481e-09 |  |  |  |  | new Film | 5787.6784
|  |  | 4.1729e-09 | 3.9119e-09 | 3.9737e-09 | 4.0462e-09 | 4.0619e-09 |  |  |  |  | Chamber | 5744.8613
|  |  | 5.6646e-09 | 5.2842e-09 | 6.1552e-09 | 4.8499e-09 | 5.5705e-09
|  |  | 7.0043e-09 | 6.5484e-09 | 7.0236e-09 | 6.6899e-09 | 6.0724e-09
|  |  | 7.9884e-09 | 7.6842e-09 | 8.1255e-09 | 7.6067e-09 | 7.8221e-09 |  |  |  |  | Fluka Exp @ 300 mAs | 6010.1887 | 5878.0764 | 5963.4634 | 5830.8083 | 6036.987 | 5906.2936 | 5228.0552 | 5903.3613
|  |  | 9.2125e-09 | 8.4823e-09 | 9.3371e-09 | 8.5597e-09 | 8.7876e-09 |  |  |  |  | Measured Exp @ 300 mAs | 5573 | 6024 | 5557 | 6013 | 5538 | 5995 | 4720 | 5935
|  |  | 9.7851e-09 | 9.5987e-09 | 1.0321e-08 | 9.4702e-09 | 9.3984e-09 |  |  |  |  | Rel. Error | 0.078448 | 0.024224 | 0.073144 | 0.0303 | 0.090102 | 0.014797 | 0.10764 | 0.0053309
|  |  | 1.0766e-08 | 1.0269e-08 | 1.0617e-08 | 1.0237e-08 | 1.0091e-08 |  |  |  |  | Correction Factor | 0.92726 | 1.0248 | 0.93184 | 1.0312 | 0.91735 | 1.015 | 0.90282 | 1.0054
|  |  | 1.104e-08 | 1.0481e-08 | 1.0841e-08 | 1.042e-08 | 1.0468e-08
|  |  | 1.6139e-08 | 1.5253e-08 | 1.665e-08 | 1.4694e-08 | 1.4818e-08
|  |  | 1.424e-08 | 1.3552e-08 | 1.4302e-08 | 1.2779e-08 | 1.3648e-08
|  |  | 2.1385e-08 | 2.0414e-08 | 2.1857e-08 | 1.9769e-08 | 1.9872e-08
|  |  | 1.3713e-08 | 1.2702e-08 | 1.3077e-08 | 1.2507e-08 | 1.253e-08
|  |  | 1.1938e-08 | 1.1366e-08 | 1.2013e-08 | 1.1285e-08 | 1.1472e-08
|  |  | 9.6007e-09 | 9.1858e-09 | 9.6039e-09 | 8.8341e-09 | 8.9545e-09
|  |  | 6.6083e-09 | 6.3543e-09 | 6.5077e-09 | 6.2995e-09 | 6.253e-09
|  |  | 3.4354e-09 | 3.3034e-09 | 3.6189e-09 | 3.21e-09 | 3.2948e-09
|  |  | 1.0732e-09 | 9.8789e-10 | 1.0289e-09 | 9.3278e-10 | 9.6209e-10
|  |  | 1.2036e-10 | 1.2062e-10 | 8.882e-11 | 1.1734e-10 | 1.1184e-10
|  |  | 3.04e-12 | 2.85e-12 | 2.01e-12 | 8.3e-13 | 3.6e-12
|  |  | 6.7e-13 | 4.9e-13 | 0 | 5e-13 | 1.63e-12
|  |  | 8.4e-13 | 5.2e-13 | 6.2e-13 | 8.3e-13 | 8.3e-13
|  |  | 1.6e-13 | 1.7e-13 | 1.2e-13 | 1.8e-13 | 2.9e-13

## InAir-WithTable
In Air Simulations for Large Collimations |  | Fluka Src: CTBOWHEEL |  |  |  | GeV/cm^2 | [5 empty] | AIR |  |  | GeV/g
With Bowtie effect and Heel effect. No Overbeaming Effect | [5 empty] | Energy FLUENCE | [5 empty] | (Linear Interpolation) |  |  | Dose
| ENERGY Bounds (GeV) | Mid-point Energy (GeV) | Mid-point Energy (KeV) | CTDI100 ¦16cm Col | QA ¦16cm Col | CTDI ¦14cm Col | QA ¦14cm Col | CTDI100 ¦12cm Col | QA ¦12cm Col | CTDI100¦8cm Col | QA ¦8cm Col | Mass Energy attenuation cm^2/g | CTDI100 ¦16cm Col | QA ¦16cm Col | CTDI100 ¦14cm Col | QA ¦14cm Col | CTDI100 ¦12cm Col | QA ¦12cm Col | CTDI100¦8cm Col | QA ¦8cm Col
1 | 0.0001505 | 0.000148 | 148 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.024892 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
2 | 0.0001455 | 0.000143 | 143 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.024721 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
3 | 0.0001405 | 0.000138 | 138 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.02455 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
4 | 0.0001355 | 0.000133 | 133 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.024379 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
5 | 0.0001305 | 0.000128 | 128 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.024208 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
6 | 0.0001255 | 0.000123 | 123 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.024037 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
7 | 0.0001205 | 0.000118 | 118 | 7.6302e-10 | 7.0007e-10 | 8.3245e-10 | 9.1058e-10 | 9.6718e-10 | 9.8691e-10 | 1.2038e-09 | 1.4834e-09 | 0.023866 | 1.821e-11 | 1.6708e-11 | 1.9867e-11 | 2.1731e-11 | 2.3082e-11 | 2.3553e-11 | 2.8728e-11 | 3.5403e-11
8 | 0.0001155 | 0.000113 | 113 | 2.3342e-09 | 2.2678e-09 | 2.633e-09 | 2.7771e-09 | 3.0341e-09 | 3.0286e-09 | 3.8249e-09 | 4.7159e-09 | 0.023695 | 5.5309e-11 | 5.3735e-11 | 6.2388e-11 | 6.5801e-11 | 7.1892e-11 | 7.1762e-11 | 9.0629e-11 | 1.1174e-10
9 | 0.0001105 | 0.000108 | 108 | 3.8046e-09 | 3.7141e-09 | 4.2832e-09 | 4.4833e-09 | 4.9657e-09 | 4.7771e-09 | 6.1729e-09 | 7.4995e-09 | 0.023524 | 8.9498e-11 | 8.737e-11 | 1.0076e-10 | 1.0546e-10 | 1.1681e-10 | 1.1237e-10 | 1.4521e-10 | 1.7641e-10
10 | 0.0001055 | 0.000103 | 103 | 5.0814e-09 | 5.0744e-09 | 5.7487e-09 | 5.937e-09 | 6.7145e-09 | 6.6018e-09 | 8.327e-09 | 1.0168e-08 | 0.023353 | 1.1866e-10 | 1.185e-10 | 1.3425e-10 | 1.3864e-10 | 1.568e-10 | 1.5417e-10 | 1.9446e-10 | 2.3744e-10
11 | 0.0001005 | 9.8e-05 | 98 | 6.2838e-09 | 6.3641e-09 | 7.0952e-09 | 7.5183e-09 | 8.3489e-09 | 8.234e-09 | 1.0342e-08 | 1.2552e-08 | 0.023332 | 1.4661e-10 | 1.4849e-10 | 1.6555e-10 | 1.7542e-10 | 1.948e-10 | 1.9211e-10 | 2.4131e-10 | 2.9286e-10
12 | 9.55e-05 | 9.3e-05 | 93 | 7.3586e-09 | 7.2015e-09 | 8.3111e-09 | 8.4762e-09 | 9.6861e-09 | 9.5431e-09 | 1.2016e-08 | 1.4633e-08 | 0.023537 | 1.732e-10 | 1.695e-10 | 1.9562e-10 | 1.995e-10 | 2.2798e-10 | 2.2462e-10 | 2.8282e-10 | 3.4442e-10
13 | 9.05e-05 | 8.8e-05 | 88 | 8.2988e-09 | 8.1559e-09 | 9.5292e-09 | 9.7438e-09 | 1.1027e-08 | 1.0782e-08 | 1.3713e-08 | 1.6688e-08 | 0.023742 | 1.9703e-10 | 1.9364e-10 | 2.2624e-10 | 2.3134e-10 | 2.618e-10 | 2.5599e-10 | 3.2556e-10 | 3.9622e-10
14 | 8.55e-05 | 8.3e-05 | 83 | 9.0858e-09 | 8.8028e-09 | 1.0385e-08 | 1.0611e-08 | 1.1967e-08 | 1.1599e-08 | 1.4806e-08 | 1.7931e-08 | 0.023947 | 2.1758e-10 | 2.108e-10 | 2.487e-10 | 2.541e-10 | 2.8657e-10 | 2.7776e-10 | 3.5455e-10 | 4.2939e-10
15 | 8.05e-05 | 7.8e-05 | 78 | 9.7263e-09 | 9.4749e-09 | 1.107e-08 | 1.1098e-08 | 1.2822e-08 | 1.2603e-08 | 1.5925e-08 | 1.9311e-08 | 0.024704 | 2.4028e-10 | 2.3407e-10 | 2.7346e-10 | 2.7417e-10 | 3.1676e-10 | 3.1135e-10 | 3.9341e-10 | 4.7705e-10
16 | 7.55e-05 | 7.3e-05 | 73 | 1.0219e-08 | 9.9855e-09 | 1.1629e-08 | 1.1843e-08 | 1.3341e-08 | 1.3005e-08 | 1.661e-08 | 2.0264e-08 | 0.026289 | 2.6865e-10 | 2.6251e-10 | 3.0572e-10 | 3.1134e-10 | 3.5073e-10 | 3.4189e-10 | 4.3665e-10 | 5.3271e-10
17 | 7.05e-05 | 6.8e-05 | 68 | 1.4917e-08 | 1.4545e-08 | 1.6769e-08 | 1.7331e-08 | 1.9473e-08 | 1.9027e-08 | 2.4208e-08 | 2.9379e-08 | 0.027874 | 4.1579e-10 | 4.0544e-10 | 4.6742e-10 | 4.831e-10 | 5.428e-10 | 5.3035e-10 | 6.7478e-10 | 8.189e-10
18 | 6.55e-05 | 6.3e-05 | 63 | 1.3189e-08 | 1.2869e-08 | 1.5027e-08 | 1.5283e-08 | 1.7391e-08 | 1.7129e-08 | 2.1621e-08 | 2.6086e-08 | 0.029459 | 3.8852e-10 | 3.791e-10 | 4.4269e-10 | 4.5023e-10 | 5.1233e-10 | 5.0459e-10 | 6.3693e-10 | 7.6848e-10 |  | CTDI100 ¦16cm Col | QA ¦16cm Col | CTDI100 ¦14cm Col | QA ¦14cm Col | CTDI100 ¦12cm Col | QA ¦12cm Col | CTDI100¦8cm Col | QA ¦8cm Col
19 | 6.05e-05 | 5.8e-05 | 58 | 1.9412e-08 | 1.9016e-08 | 2.2065e-08 | 2.2619e-08 | 2.552e-08 | 2.5006e-08 | 3.1848e-08 | 3.8628e-08 | 0.032524 | 6.3134e-10 | 6.1848e-10 | 7.1765e-10 | 7.3566e-10 | 8.3001e-10 | 8.1328e-10 | 1.0358e-09 | 1.2563e-09 | Measurement | 5573 | 6024 | 5557 | 6013 | 5538 | 5995 | 4720 | 5935
20 | 5.55e-05 | 5.3e-05 | 53 | 1.2384e-08 | 1.1995e-08 | 1.4077e-08 | 1.4423e-08 | 1.6244e-08 | 1.5723e-08 | 2.0152e-08 | 2.4364e-08 | 0.037809 | 4.6822e-10 | 4.5352e-10 | 5.3223e-10 | 5.4533e-10 | 6.1417e-10 | 5.9446e-10 | 7.6191e-10 | 9.2118e-10 | Fluka | 6010.1887 | 5878.0764 | 5963.4634 | 5830.8083 | 6036.987 | 5906.2936 | 5228.0552 | 5903.3613
21 | 5.05e-05 | 4.8e-05 | 48 | 1.1009e-08 | 1.0653e-08 | 1.2438e-08 | 1.2666e-08 | 1.4345e-08 | 1.3977e-08 | 1.7765e-08 | 2.1548e-08 | 0.04645 | 5.1135e-10 | 4.9483e-10 | 5.7774e-10 | 5.8832e-10 | 6.6632e-10 | 6.4923e-10 | 8.252e-10 | 1.0009e-09 |  | -0.078448 | 0.024224 | -0.073144 | 0.0303 | -0.090102 | 0.014797 | -0.10764 | 0.0053309
22 | 4.55e-05 | 4.3e-05 | 43 | 8.7194e-09 | 8.5527e-09 | 9.9391e-09 | 1.0214e-08 | 1.1435e-08 | 1.1062e-08 | 1.4246e-08 | 1.7278e-08 | 0.060125 | 5.2425e-10 | 5.1423e-10 | 5.9759e-10 | 6.1411e-10 | 6.8754e-10 | 6.651e-10 | 8.5654e-10 | 1.0388e-09
23 | 4.05e-05 | 3.8e-05 | 38 | 6.027e-09 | 5.7675e-09 | 6.9045e-09 | 7.0677e-09 | 7.9277e-09 | 7.823e-09 | 9.7915e-09 | 1.1894e-08 | 0.085404 | 5.1473e-10 | 4.9257e-10 | 5.8967e-10 | 6.0361e-10 | 6.7706e-10 | 6.6812e-10 | 8.3624e-10 | 1.0158e-09
24 | 3.55e-05 | 3.3e-05 | 33 | 3.1341e-09 | 3.008e-09 | 3.5809e-09 | 3.7098e-09 | 4.0906e-09 | 4.0066e-09 | 5.0788e-09 | 6.1809e-09 | 0.12809 | 4.0144e-10 | 3.8529e-10 | 4.5867e-10 | 4.7518e-10 | 5.2397e-10 | 5.1321e-10 | 6.5054e-10 | 7.9171e-10
25 | 3.05e-05 | 2.8e-05 | 28 | 9.4169e-10 | 9.0177e-10 | 1.0607e-09 | 1.1073e-09 | 1.2473e-09 | 1.1978e-09 | 1.5368e-09 | 1.8927e-09 | 0.23074 | 2.1729e-10 | 2.0807e-10 | 2.4475e-10 | 2.555e-10 | 2.8781e-10 | 2.7638e-10 | 3.5461e-10 | 4.3672e-10
26 | 2.55e-05 | 2.3e-05 | 23 | 1.0315e-10 | 1.013e-10 | 1.0673e-10 | 1.1691e-10 | 1.2949e-10 | 1.2324e-10 | 1.6206e-10 | 2.1587e-10 | 0.42334 | 4.3668e-11 | 4.2884e-11 | 4.5183e-11 | 4.9493e-11 | 5.4818e-11 | 5.2172e-11 | 6.8606e-11 | 9.1386e-11
27 | 2.05e-05 | 1.8e-05 | 18 | 1.86e-12 | 1.61e-12 | 2.02e-12 | 1.95e-12 | 1.89e-12 | 2.11e-12 | 2.56e-12 | 4.12e-12 | 0.85694 | 1.5939e-12 | 1.3797e-12 | 1.731e-12 | 1.671e-12 | 1.6196e-12 | 1.8081e-12 | 2.1938e-12 | 3.5306e-12
28 | 1.55e-05 | 1.3e-05 | 13 | 5.8e-13 | 1.03e-12 | 6.3e-13 | 6.3e-13 | 6.2e-13 | 4.2e-13 | 1.08e-12 | 1.13e-12 | 2.6972 | 1.5644e-12 | 2.7781e-12 | 1.6992e-12 | 1.6992e-12 | 1.6723e-12 | 1.1328e-12 | 2.913e-12 | 3.0478e-12
29 | 1.05e-05 | 8e-06 | 8 | 4e-13 | 5.9e-13 | 4.5e-13 | 4.5e-13 | 3.8e-13 | 6.8e-13 | 7.4e-13 | 7.1e-13 | 9.446 | 3.7784e-12 | 5.5731e-12 | 4.2507e-12 | 4.2507e-12 | 3.5895e-12 | 6.4233e-12 | 6.99e-12 | 6.7067e-12
30 | 5.5e-06 | 3e-06 | 3 | 1.1e-13 | 1.4e-13 | 1.7e-13 | 1e-13 | 1.5e-13 | 1.4e-13 | 2.2e-13 | 2e-13 | 161.4 | 1.7754e-11 | 2.2596e-11 | 2.7438e-11 | 1.614e-11 | 2.421e-11 | 2.2596e-11 | 3.5508e-11 | 3.228e-11
| 5e-07 |  |  |  |  | 1.7349e-07 | 1.7794e-07 | 2.0068e-07 | 1.9624e-07 | 2.4935e-07 | 3.0272e-07 | 0.36048 | 5.6663e-09 | 5.5221e-09 | 6.4412e-09 | 6.6018e-09 | 7.4351e-09 | 7.2644e-09 | 9.2421e-09 | 1.1219e-08
[12 empty] | Dose (mGy/particle) | 9.0784e-13 | 8.8473e-13 | 1.032e-12 | 1.0577e-12 | 1.1912e-12 | 1.1639e-12 | 1.4808e-12 | 1.7976e-12
| Constants | [10 empty] | Exposure mR /particle | 1.0364e-10 | 1.01e-10 | 1.1781e-10 | 1.2075e-10 | 1.3599e-10 | 1.3286e-10 | 1.6904e-10 | 2.052e-10
[7 empty] | 3.22e-07 | 2e-07
1.61
[12 empty] | Fluka Exp @ 300 mAs | 5756.1462 | 5609.6005 | 5757.845 | 5901.3736 | 5841.6688 | 5707.5612 | 4959.407 | 6020.4715
[11 empty] | (Measured at 1s Rotation) | Measured Exp @ 300 mAs | 5573 | 6024 | 5557 | 6013 | 5538 | 5995 | 4720 | 5935
[12 empty] | Rel. Error | 0.032863 | 0.068791 | 0.036143 | 0.018564 | 0.054834 | 0.047946 | 0.050722 | 0.014401
Source Strength @ 300 mAs | [11 empty] | Correction Factor | 0.96818 | 1.0739 | 0.96512 | 1.0189 | 0.94802 | 1.0504 | 0.95173 | 0.9858
Gev/G to mGy | 0.00016022 |  |  |  | Expected Src Strengths
Collimation (cm) | Bowtie | Bowtie+FilmHEEL | Bowtie+ChamberHEEL | Bowtie+newFilm | CTDI | QA | Mean |  |  |  |  | Old-Film | Day 2
16 | 5.6676e+13 | 5.5542e+13 | 5.3411e+13 | 5.0733e+13 | 5.3775e+13 | 5.9646e+13 |  |  |  |  | Collimation | Fluka 1.9cm chamber | Measured 1.9cm Chamber
14 | 5.021e+13 | 48874646514455 |  |  | 4.717e+13 | 1.0189 |  |  |  |  | 16 | 5609.6005 | 6024
12 | 43345505620919 | 4.2958e+13 |  |  | 0.94802 | 4.5121e+13 |  |  |  |  | 14 | 5901.3736 | 6013
8 | 2.9669e+13 | 2.9339e+13 |  |  | 2.7923e+13 | 2.8923e+13 |  |  |  |  | 12 | 5707.5612 | 5995 | 1.0329 |  | Average Correction Factor (CTDI Chamber)
[11 empty] | 8 | 6020.4715 | 5935 |  |  | 0.95826
160 Collimation
|  |  | CTDI160 NEW Film | CTDI160 OlD Film | QA160 New Film | QA160 Old Film | QA chamber
| Day2 Correction Factor for source strength |  | 0 | 0 | 0 | 0 | 0
| 1.008 |  | 0 | 0 | 0 | 0 | 0 |  |  |  |  | 160 Collimation - QA chamber Measurement
|  |  | 0 | 0 | 0 | 0 | 0 |  |  |  |  | Day 1 | 5978
|  |  | 0 | 0 | 0 | 0 | 0 |  |  |  |  | Day 2 | 6024
|  |  | 0 | 0 | 0 | 0 | 0
|  |  | 0 | 0 | 0 | 0 | 0 |  |  |  |  | 160 Collimation - Heel differences - Day 1 Source Strength
|  |  | 8.2089e-10 | 8.2027e-10 | 8.985e-10 | 8.1692e-10 | 7.1086e-10 |  |  |  |  | old Film | 5878.0764
|  |  | 2.5839e-09 | 2.4272e-09 | 2.538e-09 | 2.3726e-09 | 2.5481e-09 |  |  |  |  | new Film | 5787.6784
|  |  | 4.1729e-09 | 3.9119e-09 | 3.9737e-09 | 4.0462e-09 | 4.0619e-09 |  |  |  |  | Chamber | 5744.8613
|  |  | 5.6646e-09 | 5.2842e-09 | 6.1552e-09 | 4.8499e-09 | 5.5705e-09
|  |  | 7.0043e-09 | 6.5484e-09 | 7.0236e-09 | 6.6899e-09 | 6.0724e-09
|  |  | 7.9884e-09 | 7.6842e-09 | 8.1255e-09 | 7.6067e-09 | 7.8221e-09 |  |  |  |  | Fluka Exp @ 300 mAs | 5756.1462 | 5609.6005 | 5757.845 | 5901.3736 | 5841.6688 | 5707.5612 | 4959.407 | 6020.4715
|  |  | 9.2125e-09 | 8.4823e-09 | 9.3371e-09 | 8.5597e-09 | 8.7876e-09 |  |  |  |  | Measured Exp @ 300 mAs | 5573 | 6024 | 5557 | 6013 | 5538 | 5995 | 4720 | 5935
|  |  | 9.7851e-09 | 9.5987e-09 | 1.0321e-08 | 9.4702e-09 | 9.3984e-09 |  |  |  |  | Rel. Error | 0.032863 | 0.068791 | 0.036143 | 0.018564 | 0.054834 | 0.047946 | 0.050722 | 0.014401
|  |  | 1.0766e-08 | 1.0269e-08 | 1.0617e-08 | 1.0237e-08 | 1.0091e-08 |  |  |  |  | Correction Factor | 0.96818 | 1.0739 | 0.96512 | 1.0189 | 0.94802 | 1.0504 | 0.95173 | 0.9858
|  |  | 1.104e-08 | 1.0481e-08 | 1.0841e-08 | 1.042e-08 | 1.0468e-08
|  |  | 1.6139e-08 | 1.5253e-08 | 1.665e-08 | 1.4694e-08 | 1.4818e-08
|  |  | 1.424e-08 | 1.3552e-08 | 1.4302e-08 | 1.2779e-08 | 1.3648e-08
|  |  | 2.1385e-08 | 2.0414e-08 | 2.1857e-08 | 1.9769e-08 | 1.9872e-08
|  |  | 1.3713e-08 | 1.2702e-08 | 1.3077e-08 | 1.2507e-08 | 1.253e-08
|  |  | 1.1938e-08 | 1.1366e-08 | 1.2013e-08 | 1.1285e-08 | 1.1472e-08
|  |  | 9.6007e-09 | 9.1858e-09 | 9.6039e-09 | 8.8341e-09 | 8.9545e-09
|  |  | 6.6083e-09 | 6.3543e-09 | 6.5077e-09 | 6.2995e-09 | 6.253e-09
|  |  | 3.4354e-09 | 3.3034e-09 | 3.6189e-09 | 3.21e-09 | 3.2948e-09
|  |  | 1.0732e-09 | 9.8789e-10 | 1.0289e-09 | 9.3278e-10 | 9.6209e-10
|  |  | 1.2036e-10 | 1.2062e-10 | 8.882e-11 | 1.1734e-10 | 1.1184e-10
|  |  | 3.04e-12 | 2.85e-12 | 2.01e-12 | 8.3e-13 | 3.6e-12
|  |  | 6.7e-13 | 4.9e-13 | 0 | 5e-13 | 1.63e-12
|  |  | 8.4e-13 | 5.2e-13 | 6.2e-13 | 8.3e-13 | 8.3e-13
|  |  | 1.6e-13 | 1.7e-13 | 1.2e-13 | 1.8e-13 | 2.9e-13

## CTDI
In Air Simulations for Large Collimations |  | Fluka Src: CTBOWHEEL |  |  |  | GeV/cm^2 | [5 empty] | AIR |  |  | GeV/g
With Bowtie effect and Heel effect. No Overbeaming Effect | [5 empty] | Energy FLUENCE | [5 empty] | (Linear Interpolation) |  |  | Dose
| ENERGY Bounds (GeV) | Mid-point Energy (GeV) | Mid-point Energy (KeV) | A ¦16cm Col | B ¦16cm Col | A ¦14cm Col | B ¦14cm Col | A ¦12cm Col | B ¦12cm Col | A¦80cm Col | B ¦80cm Col | Mass Energy attenuation cm^2/g | A ¦16cm Col | B ¦16cm Col | A ¦14cm Col | B ¦14cm Col | A ¦12cm Col | B ¦12cm Col | A¦80cm Col | B ¦80cm Col
1 | 0.0001505 | 0.000148 | 148 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.024892 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
2 | 0.0001455 | 0.000143 | 143 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.024721 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
3 | 0.0001405 | 0.000138 | 138 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.02455 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
4 | 0.0001355 | 0.000133 | 133 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.024379 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
5 | 0.0001305 | 0.000128 | 128 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.024208 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
6 | 0.0001255 | 0.000123 | 123 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.024037 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
7 | 0.0001205 | 0.000118 | 118 | 4.7906e-11 | 2.6868e-10 | 5.635e-11 | 2.6731e-10 | 6.5007e-11 | 3.4411e-10 | 7.4068e-11 | 3.6873e-10 | 0.023866 | 1.1433e-12 | 6.4122e-12 | 1.3448e-12 | 6.3795e-12 | 1.5514e-12 | 8.2125e-12 | 1.7677e-12 | 8.8e-12
8 | 0.0001155 | 0.000113 | 113 | 1.8128e-10 | 7.8988e-10 | 2.3209e-10 | 9.5588e-10 | 2.259e-10 | 1.1105e-09 | 3.0601e-10 | 1.1379e-09 | 0.023695 | 4.2954e-12 | 1.8716e-11 | 5.4993e-12 | 2.2649e-11 | 5.3526e-12 | 2.6312e-11 | 7.2507e-12 | 2.6962e-11
9 | 0.0001105 | 0.000108 | 108 | 3.4386e-10 | 1.4834e-09 | 3.7859e-10 | 1.5835e-09 | 4.255e-10 | 1.6931e-09 | 5.4831e-10 | 1.8785e-09 | 0.023524 | 8.0889e-12 | 3.4895e-11 | 8.9057e-12 | 3.725e-11 | 1.0009e-11 | 3.9827e-11 | 1.2898e-11 | 4.4189e-11
10 | 0.0001055 | 0.000103 | 103 | 5.1786e-10 | 1.9979e-09 | 6.2351e-10 | 2.1585e-09 | 7.1146e-10 | 2.4795e-09 | 7.8098e-10 | 2.6736e-09 | 0.023353 | 1.2093e-11 | 4.6657e-11 | 1.4561e-11 | 5.0407e-11 | 1.6614e-11 | 5.7904e-11 | 1.8238e-11 | 6.2435e-11
11 | 0.0001005 | 9.8e-05 | 98 | 7.1714e-10 | 2.4389e-09 | 8.2034e-10 | 2.8126e-09 | 9.0719e-10 | 2.9776e-09 | 1.0315e-09 | 3.3523e-09 | 0.023332 | 1.6732e-11 | 5.6904e-11 | 1.914e-11 | 6.5624e-11 | 2.1167e-11 | 6.9473e-11 | 2.4067e-11 | 7.8215e-11
12 | 9.55e-05 | 9.3e-05 | 93 | 9.0902e-10 | 2.9557e-09 | 9.1935e-10 | 3.488e-09 | 1.0355e-09 | 3.8614e-09 | 1.2975e-09 | 4.169e-09 | 0.023537 | 2.1396e-11 | 6.9569e-11 | 2.1639e-11 | 8.2096e-11 | 2.4374e-11 | 9.0885e-11 | 3.0538e-11 | 9.8126e-11
13 | 9.05e-05 | 8.8e-05 | 88 | 1.1533e-09 | 3.4477e-09 | 1.1692e-09 | 4.1373e-09 | 1.3313e-09 | 4.4167e-09 | 1.6239e-09 | 4.8125e-09 | 0.023742 | 2.7381e-11 | 8.1856e-11 | 2.776e-11 | 9.8227e-11 | 3.1607e-11 | 1.0486e-10 | 3.8554e-11 | 1.1426e-10
14 | 8.55e-05 | 8.3e-05 | 83 | 1.3386e-09 | 4.1564e-09 | 1.4474e-09 | 4.584e-09 | 1.6605e-09 | 5.0437e-09 | 1.8636e-09 | 5.4804e-09 | 0.023947 | 3.2056e-11 | 9.9532e-11 | 3.4662e-11 | 1.0977e-10 | 3.9765e-11 | 1.2078e-10 | 4.4628e-11 | 1.3124e-10
15 | 8.05e-05 | 7.8e-05 | 78 | 1.5869e-09 | 4.7555e-09 | 1.6973e-09 | 5.2336e-09 | 1.9004e-09 | 5.8796e-09 | 2.2571e-09 | 6.194e-09 | 0.024704 | 3.9202e-11 | 1.1748e-10 | 4.193e-11 | 1.2929e-10 | 4.6946e-11 | 1.4525e-10 | 5.576e-11 | 1.5302e-10
16 | 7.55e-05 | 7.3e-05 | 73 | 1.8376e-09 | 5.3643e-09 | 2.0295e-09 | 5.922e-09 | 2.1828e-09 | 6.4155e-09 | 2.4818e-09 | 6.9557e-09 | 0.026289 | 4.8309e-11 | 1.4102e-10 | 5.3352e-11 | 1.5568e-10 | 5.7384e-11 | 1.6866e-10 | 6.5243e-11 | 1.8286e-10
17 | 7.05e-05 | 6.8e-05 | 68 | 2.3626e-09 | 7.102e-09 | 2.6182e-09 | 7.8847e-09 | 2.7555e-09 | 8.5445e-09 | 3.2798e-09 | 9.3851e-09 | 0.027874 | 6.5856e-11 | 1.9796e-10 | 7.2979e-11 | 2.1978e-10 | 7.6806e-11 | 2.3817e-10 | 9.1422e-11 | 2.616e-10
18 | 6.55e-05 | 6.3e-05 | 63 | 2.6322e-09 | 7.0461e-09 | 2.8181e-09 | 7.9114e-09 | 3.1573e-09 | 8.6294e-09 | 3.611e-09 | 9.263e-09 | 0.029459 | 7.7542e-11 | 2.0757e-10 | 8.3017e-11 | 2.3306e-10 | 9.3012e-11 | 2.5421e-10 | 1.0638e-10 | 2.7288e-10
19 | 6.05e-05 | 5.8e-05 | 58 | 3.4169e-09 | 9.6234e-09 | 3.6811e-09 | 1.0603e-08 | 4.1336e-09 | 1.1623e-08 | 4.5877e-09 | 1.2577e-08 | 0.032524 | 1.1113e-10 | 3.1299e-10 | 1.1972e-10 | 3.4486e-10 | 1.3444e-10 | 3.7802e-10 | 1.4921e-10 | 4.0904e-10
20 | 5.55e-05 | 5.3e-05 | 53 | 3.3937e-09 | 8.0071e-09 | 3.6403e-09 | 8.6701e-09 | 3.9364e-09 | 9.2729e-09 | 4.4402e-09 | 1.0052e-08 | 0.037809 | 1.2831e-10 | 3.0274e-10 | 1.3764e-10 | 3.2781e-10 | 1.4883e-10 | 3.506e-10 | 1.6788e-10 | 3.8006e-10
21 | 5.05e-05 | 4.8e-05 | 48 | 3.475e-09 | 7.4111e-09 | 3.8354e-09 | 8.1972e-09 | 3.9978e-09 | 8.7447e-09 | 4.5272e-09 | 9.4342e-09 | 0.04645 | 1.6141e-10 | 3.4425e-10 | 1.7815e-10 | 3.8076e-10 | 1.857e-10 | 4.0619e-10 | 2.1029e-10 | 4.3822e-10
22 | 4.55e-05 | 4.3e-05 | 43 | 3.2449e-09 | 5.9785e-09 | 3.5143e-09 | 6.5384e-09 | 3.7476e-09 | 7.0993e-09 | 4.234e-09 | 7.8852e-09 | 0.060125 | 1.951e-10 | 3.5946e-10 | 2.1129e-10 | 3.9312e-10 | 2.2533e-10 | 4.2684e-10 | 2.5457e-10 | 4.741e-10
23 | 4.05e-05 | 3.8e-05 | 38 | 2.6529e-09 | 4.4165e-09 | 2.8438e-09 | 4.5744e-09 | 3.0384e-09 | 5.0621e-09 | 3.3542e-09 | 5.4573e-09 | 0.085404 | 2.2657e-10 | 3.7719e-10 | 2.4287e-10 | 3.9068e-10 | 2.5949e-10 | 4.3232e-10 | 2.8646e-10 | 4.6608e-10
24 | 3.55e-05 | 3.3e-05 | 33 | 1.6457e-09 | 2.2928e-09 | 1.7079e-09 | 2.4987e-09 | 1.8348e-09 | 2.7319e-09 | 2.0537e-09 | 2.9005e-09 | 0.12809 | 2.108e-10 | 2.9368e-10 | 2.1876e-10 | 3.2006e-10 | 2.3502e-10 | 3.4993e-10 | 2.6306e-10 | 3.7152e-10
25 | 3.05e-05 | 2.8e-05 | 28 | 6.1511e-10 | 7.8042e-10 | 6.5616e-10 | 8.1663e-10 | 7.0199e-10 | 8.6415e-10 | 7.8214e-10 | 9.6716e-10 | 0.23074 | 1.4193e-10 | 1.8008e-10 | 1.514e-10 | 1.8843e-10 | 1.6198e-10 | 1.9939e-10 | 1.8047e-10 | 2.2316e-10
26 | 2.55e-05 | 2.3e-05 | 23 | 8.572e-11 | 9.409e-11 | 9.133e-11 | 1.0707e-10 | 9.6758e-11 | 9.8744e-11 | 9.1937e-11 | 1.1621e-10 | 0.42334 | 3.6289e-11 | 3.9832e-11 | 3.8664e-11 | 4.5328e-11 | 4.0962e-11 | 4.1802e-11 | 3.8921e-11 | 4.9196e-11
27 | 2.05e-05 | 1.8e-05 | 18 | 2.167e-12 | 1.471e-12 | 1.858e-12 | 1.414e-12 | 2.547e-12 | 1.634e-12 | 2.434e-12 | 3.1e-12 | 0.85694 | 1.857e-12 | 1.2606e-12 | 1.5922e-12 | 1.2117e-12 | 2.1826e-12 | 1.4002e-12 | 2.0858e-12 | 2.6565e-12
28 | 1.55e-05 | 1.3e-05 | 13 | 1.49e-13 | 6.4e-14 | 0 | 0 | 4.2e-13 | 5.69e-13 | 5e-13 | 0 | 2.6972 | 4.0188e-13 | 1.7262e-13 | 0 | 0 | 1.1328e-12 | 1.5347e-12 | 1.3486e-12 | 0
29 | 1.05e-05 | 8e-06 | 8 | 1.11e-13 | 8.5e-14 | 4e-15 | 4.3e-14 | 7.8e-14 | 1.28e-13 | 1.43e-13 | 5e-14 | 9.446 | 1.0485e-12 | 8.0291e-13 | 3.7784e-14 | 4.0618e-13 | 7.3679e-13 | 1.2091e-12 | 1.3508e-12 | 4.723e-13 |  |  | 0.30902
30 | 5.5e-06 | 3e-06 | 3 | 4.2e-14 | 5.7e-14 | 0 | 5.7e-14 | 3.9e-14 | 7.8e-14 | 3.5e-14 | 5e-14 | 161.4 | 6.7788e-12 | 9.1998e-12 | 0 | 9.1998e-12 | 6.2946e-12 | 1.2589e-11 | 5.649e-12 | 8.07e-12
| 5e-07 |  |  |  |  | 3.4782e-08 | 8.8946e-08 | 3.7849e-08 | 9.6895e-08 | 4.323e-08 | 1.0506e-07 | 0.36048 | 1.5757e-09 | 3.3002e-09 | 1.6849e-09 | 3.6121e-09 | 1.8267e-09 | 3.9264e-09 | 2.058e-09 | 4.2572e-09
[12 empty] | Dose (mGy/particle) | 2.5246e-13 | 5.2875e-13 | 2.6996e-13 | 5.7872e-13 | 2.9267e-13 | 6.2908e-13 | 3.2973e-13 | 6.8207e-13
| Constants | [10 empty] | Exposure mR /particle | 2.882e-11 | 6.036e-11 | 3.0817e-11 | 6.6064e-11 | 3.3409e-11 | 7.1812e-11 | 3.7641e-11 | 7.7862e-11
[7 empty] | 3.22e-07 | 2e-07
[7 empty] | 1.61 | [5 empty] | 2.5246e-13 | 2.4834e-13 | 2.7226e-13
2284.4319
[12 empty] | Fluka Exp @ 300 mAs | 1600.7109 | 3352.5382 | 1506.1604 | 3228.8422 | 1435.1937 | 3084.8965 | 1104.3626 | 2284.4319
[11 empty] | (Measured at 1s Rotation) | Measured Exp @ 300 mAs | 1391 |  | 1330 |  | 1237 |  | 948
[12 empty] | Rel. Error | 0.15076 | #DIV/0! | 0.13245 | #DIV/0! | 0.16022 | #DIV/0! | 0.16494 | #DIV/0! | Mean Phantom CF
| Simulation FILES Folder : InAirRuns | [10 empty] | Corr Fac. Phantom meas | 0.86899 |  | 0.88304 |  | 0.8619 |  | 0.85841
[12 empty] | Measured Exp @ 300mAs [Static] | 1454 |  | 1392 |  | 1295 |  | 993
Source Strength @ 300 mAs
Gev/G to mGy | 0.00016022 |  |  |  | Expected Src Strengths
Collimation (cm) | Bowtie | Bowtie+FilmHEEL | Bowtie+ChamberHEEL | Bowtie+NewFilm | CTDI | QA | Mean
16 | 5.6676e+13 | 5.5542e+13 | 5.3411e+13 | 5.0733e+13 | 4.8266e+13 | 0 | [8 empty] | Correction Factors for CTDI100 chamber
14 | 5.021e+13 | 48874646514455 |  |  | 4.3158e+13 | 0 | [8 empty] | 0.87 |  |  | PLOTTING
12 | 43345505620919 | 4.2958e+13 |  |  | 0.8619 | 0 | [5 empty] | Collimation | Fluka Exp @ 300 mAs | Measured Exp @ 300 mAs | Fluka Exp (Corrected) | Rel.Error |  |  | Measurement | Fluka | Fluka Corrected | CF_CTDI_Measurement
8 | 2.9669e+13 | 2.9339e+13 |  |  | 2.5185e+13 | 0 | [5 empty] | 16 | 1604.1449 | 1391 | 1395.6061 | 0.0033113 |  | 16cm Col | 1391 | 1604.1449 | 1395.6061 | 0.86713
[12 empty] | 14 | 1506.1604 | 1330 | 1310.3596 | 0.014767 |  | 14cm Col | 1330 | 1506.1604 | 1310.3596 | 0.88304
[12 empty] | 12 | 1435.1937 | 1237 | 1248.6185 | 0.0093925 |  | 12cm Col | 1237 | 1435.1937 | 1248.6185 | 0.8619
|  |  | Old Film - 2E8 | Old Film 1E7 |  |  | Chamber |  |  | no Heel |  | 8 | 1104.3626 | 948 | 960.7955 | 0.013497 |  | 8cm Col | 948 | 1104.3626 | 960.7955 | 0.85841
|  |  | CTDI16-A | CTDI16-A | CTDI-B |  | CTDI16-A | CTDI-B |  | CTDI16-A | [11 empty] | 0.86762
|  |  | 0 | 0 | 0 |  | 0 | 0 |  | 0
|  |  | 0 | 0 | 0 |  | 0 | 0 |  | 0 |  |  |  |  | 7.274%
|  |  | 0 | 0 | 0 |  | 0 | 0 |  | 0 | [7 empty] | 7.844%
|  |  | 0 | 0 | 0 |  | 0 | 0 |  | 0 | [7 empty] | 7.274%
|  |  | 0 | 0 | 0 |  | 0 | 0 |  | 0
|  |  | 0 | 0 | 0 |  | 0 | 0 |  | 0
|  |  | 4.7906e-11 | 5.3917e-11 | 2.6868e-10 |  | 3.4025e-11 | 2.5896e-10 |  | 3.9523e-11
|  |  | 1.8128e-10 | 1.9259e-10 | 7.8988e-10 |  | 1.4143e-10 | 7.9659e-10 |  | 1.9222e-10
|  |  | 3.4386e-10 | 3.3279e-10 | 1.4834e-09 |  | 3.7209e-10 | 1.4847e-09 |  | 3.8616e-10
|  |  | 5.1786e-10 | 5.3796e-10 | 1.9979e-09 |  | 5.4072e-10 | 1.9706e-09 |  | 5.2136e-10
|  |  | 7.1714e-10 | 7.1144e-10 | 2.4389e-09 |  | 7.9451e-10 | 2.5109e-09 |  | 7.2738e-10
|  |  | 9.0902e-10 | 8.7494e-10 | 2.9557e-09 |  | 1.0057e-09 | 3.0688e-09 |  | 8.7332e-10
|  |  | 1.1533e-09 | 1.1465e-09 | 3.4477e-09 |  | 1.1552e-09 | 3.5967e-09 |  | 1.2041e-09
|  |  | 1.3386e-09 | 1.3707e-09 | 4.1564e-09 |  | 1.4571e-09 | 4.2568e-09 |  | 1.3349e-09
|  |  | 1.5869e-09 | 1.6131e-09 | 4.7555e-09 |  | 1.5815e-09 | 4.8034e-09 |  | 1.7191e-09
|  |  | 1.8376e-09 | 1.8446e-09 | 5.3643e-09 |  | 1.8853e-09 | 5.4533e-09 |  | 1.8978e-09
|  |  | 2.3626e-09 | 2.3362e-09 | 7.102e-09 |  | 2.3456e-09 | 7.1894e-09 |  | 2.4866e-09
|  |  | 2.6322e-09 | 2.6615e-09 | 7.0461e-09 |  | 2.7563e-09 | 7.1963e-09 |  | 2.7235e-09
|  |  | 3.4169e-09 | 3.4377e-09 | 9.6234e-09 |  | 3.4925e-09 | 9.4911e-09 |  | 3.6508e-09
|  |  | 3.3937e-09 | 3.3726e-09 | 8.0071e-09 |  | 3.472e-09 | 8.0165e-09 |  | 3.4324e-09
|  |  | 3.475e-09 | 3.4576e-09 | 7.4111e-09 |  | 3.6192e-09 | 7.3675e-09 |  | 3.4512e-09
|  |  | 3.2449e-09 | 3.2568e-09 | 5.9785e-09 |  | 3.3123e-09 | 6.1427e-09 |  | 3.3795e-09
|  |  | 2.6529e-09 | 2.6869e-09 | 4.4165e-09 |  | 2.5621e-09 | 4.2634e-09 |  | 2.66e-09
|  |  | 1.6457e-09 | 1.6534e-09 | 2.2928e-09 |  | 1.6436e-09 | 2.3223e-09 |  | 1.6778e-09
|  |  | 6.1511e-10 | 6.1243e-10 | 7.8042e-10 |  | 6.5165e-10 | 7.1548e-10 |  | 6.4597e-10
|  |  | 8.572e-11 | 8.8854e-11 | 9.409e-11 |  | 9.0765e-11 | 9.4303e-11 |  | 9.8343e-11
|  |  | 2.167e-12 | 1.101e-12 | 1.471e-12 |  | 1.478e-12 | 2.011e-12 |  | 1.215e-12
|  |  | 1.49e-13 | 1.03e-13 | 6.4e-14 |  | 3.2e-14 | 1.92e-13 |  | 6.22e-13
|  |  | 1.11e-13 | 8.5e-14 | 8.5e-14 |  | 2.1e-14 | 1.63e-13 |  | 3e-15
|  |  | 4.2e-14 | 3.6e-14 | 5.7e-14 |  | 2.1e-14 | 9.2e-14 |  | 4e-15

## 5yrOLD
MOSFET-ID | Phantom Location |  | Slice #
Reader A | Organ | Hole |  | Average Measured Dose (cGy) | Fluka Dose (cGy) (carbon table+ preproc Vxl) | MATLAB | Fluka Dose (cGy) No Preporcessing | EyeBalling | X (cm) | Y (cm) | Z (cm)
A-1 | Lung | 81 | 14 | 1.9675 | 3.7073 | 3.95 |  |  | 3.209 | 1.747 | 4.5
A-2 | Lung | 84 | 14 | 2.2839 | 3.4491 | 3.6 |  |  | 5.809 | -2.112 | 4.5
A-3 | Lung | 77 | 14 | 2.1842 | 3.561 | 3.7414 |  |  | -5.972 | 3.209 | 4.5
A-4 | Lung | 79 | 14 | 1.964 | 3.2211 | 3.44 |  |  | -5.119 | -3.291 | 4.5
A-5 | ST | 85 | 14 | 1.6984 | 2.8936 | 2.8 |  |  | -0.041 | 0.651 | 4.5
B-1 | Lung | 68 | 13 | 1.9812 | 3.8275 |  |  |  | -6.132 | -0.325 | 7
B-2 | Lung | 72 | 13 | 2.5282 | 4.0242 |  |  |  | 6.338 | 1.138 | 7
B-3 | Bone | 63 | 13 | 5.5442 | 10.4037 | 7.81 |  |  | -4.712 | -4.591 | 7
B-4 | Bone | 65 | 13 | 5.2472 | 10.3538 | 8.61 |  |  | 7.8 | -2.031 | 7
B-5 | Bone | 76 | 13 | 6.7322 | 10.2577 | 7.78 |  |  | -0.325 | 7.272 | 7
C-1 | ST | 58 | 12 | 2.7937 | 3.699 |  |  |  | -6.094 | 5.85 | 9.5
C-2 | ST | 59 | 12 | 2.7937 | 3.6352 |  |  |  | 5.769 | 6.175 | 9.5
C-3 | ST | 60 | 12 | 2.7874 | 3.688 |  |  |  | -0.041 | 4.712 | 9.5
C-4 | ST/Bone | 51 | 12 | 2.054 | 3.0568 |  |  |  | -0.488 | -0.772 | 9.5
C-5 | ST | 50 | 11 | 2.6223 | 3.7464 |  |  |  | -0.041 | 4.225 | 12
D-1 | Bone | 39 | 11 | 11.3743 | 11.8578 |  |  |  | -0.041 | 5.85 | 12
D-2 | Bone | 37 | 11 | 9.8673 | 10.7543 |  |  |  | 7.434 | -2.478 | 12
D-3 | Bone | 38 | 11 | 7.4692 | 9.0457 |  |  |  | -0.203 | -2.234 | 12
D-4 | ST/Bone | 44 | 11 | 3.2731 | 3.3391 |  |  |  | 7.841 | 0.488 | 12
D-5 | Lung | 48 | 11 | 3.5979 | 4.1625 |  |  |  | 3.819 | 2.234 | 12
E-1 | Bone | 30 | 10 | 9.7793 | 10.2506 |  |  |  | -8.084 | 0.569 | 14.5 | Dose in centi Gray
E-2 | Bone | 33 | 10 | 8.8333 | 10.8913 |  |  |  | 8.328 | 0.2 | 14.5 | Hole | Measured | Carbon Table | No Table
E-3 | ST | 34 | 10 | 3.108 | 3.4965 |  |  |  | -0.041 | 3.372 | 14.5 | 50 | 2.6223 | 3.7 | 3.814
E-4 | ST | 35 | 10 | 2.6667 | 3.0288 |  |  |  | -0.244 | -0.284 | 14.5 | 39 | 11.3743 | 11.2 | 12.3143
E-5 | Lung | 28 | 10 | 3.2677 | 3.6643 |  |  |  | -3.819 | -0.569 | 14.5 | 37 | 9.8673 | 9.3 | 11.9921
[12 empty] | 38 | 7.4692 | 8 | 10.2284
[12 empty] | 44 | 3.2731 | 3.3 | 3.5559
[12 empty] | 48 | 3.5979 | 4 | 4.6942
Slab 11
Slab 13
Slab 12
Slab 14
Slab 10

## 5 yr old - multiple runs
[9 empty] | 5FT.vxl VXL, with Carbon density 1.2g/cc and Foam |  |  |  | Voxel File VXL, with Carbon density 1.2g/cc and Foam |  | Full Carbon Table (No Foam assumed) |  | Voxel File VXL, with Carbon density 1.2g/cc and Foam
|  |  |  | Average Measured Dose (cGy) |  |  |  |  | Fluka Simlations |  | Additional Simulations/Modifications |  | Fluka Dose (Head First) - Chamber Heel |  | Old FILM |  | Fluka Dose (Feet First) - Chamber Heel
Organ | Hole | Organ-Hole | Slab | Run 1 (Head First) - Old Calibration | Run 2 (Head First) - Old Calibration | Run 3 ( Feet First) - Old Calibration | Run 2 (Head First) - New Calibration | Run 3 (Feet First) - New Calibration | Fluka Dose (Head First) | Fluka Dose (Feet First) |  | Feet First (Corrected Fluka Dose) | No Correction | With Correction |  | Fluka Dose (Head First) | No Correction | With Correction
Lung | 81 | Lung-81 | 14 | 1.9675 | 2.8107 | 3.7026 | 2.8379 | 3.7385 | 3.9047 | 4.3269 |  | 3.9807 | 3.4368 | 3.1619 |  | 3.7073 | 4.7449 | 4.3653
Lung | 84 | Lung-84 | 14 | 2.2839 | 2.282 | 3.0488 | 2.6693 | 3.5663 | 3.8169 | 4.1921 |  | 3.8567 | 3.1165 | 2.8672 |  | 3.4491 | 4.6622 | 4.2892
Lung | 77 | Lung-77 | 14 | 2.1842 | 2.4585 | 3.1878 | 2.6642 | 3.4546 | 4.0202 | 4.3576 |  | 4.009 | 2.8341 | 2.6074 |  | 3.561 | 4.2678 | 3.9263
Lung | 79 | Lung-79 | 14 | 1.964 | 2.1006 | 2.7452 | 2.3002 | 3.006 | 3.6274 | 4.1133 |  | 3.7843 | 2.8697 | 2.6402 |  | 3.2211 | 4.4523 | 4.0961
ST | 85 | ST-85 | 14 | 1.6984 | 2.3295 | 3.0908 | 2.6912 | 3.5707 | 3.1653 | 3.4719 |  | 3.1941 | 2.4453 | 2.2497 |  | 2.8936 | 3.6943 | 3.3988
Lung | 68 | Lung-68 | 13 | 1.9812 | 2.9975 | 3.5456 | 3.6047 | 4.2639 | 3.9757 | 4.5181 |  | 4.1567 | 3.5419 | 3.2586 |  | 3.8275 | 4.6912 | 4.3159
Lung | 72 | Lung-72 | 13 | 2.5282 | 3.1242 | 3.6218 | 3.8641 | 4.4796 | 4.2725 | 4.461 |  | 4.1041 | 3.9072 | 3.5947 |  | 4.0242 | 4.6308 | 4.2604
Bone | 63 | Bone-63 | 13 | 5.5442 | 6.8688 | 8.3165 | 9.0031 | 10.9006 | 10.3503 | 11.007 |  | 10.1265 | 9.5894 | 8.8223 |  | 10.4037 | 11.2992 | 10.3953
Bone | 65 | Bone-65 | 13 | 5.2472 | 9.2853 | 10.2797 | 10.1508 | 11.2379 | 9.8377 | 10.7347 |  | 9.8759 | 8.9391 | 8.224 |  | 10.3538 | 11.0776 | 10.1914
Bone | 76 | Bone-76 | 13 | 6.7322 | 10.1522 | 11.8659 | 10.8915 | 12.7299 | 11.4368 | 12.3908 |  | 11.3995 | 10.4623 | 9.6253 |  | 10.2577 | 12.8755 | 11.8455
ST | 58 | ST-58 | 12 | 2.7937 | 2.7425 | 2.7951 | 2.8833 | 2.9386 | 3.6814 | 3.7793 |  | 3.4769 | 3.6081 | 3.3194 |  | 3.699 | 3.939 | 3.6239
ST | 59 | ST-59 | 12 | 2.7937 | 2.7326 | 2.8737 | 2.7262 | 2.867 | 3.724 | 3.6933 |  | 3.3978 | 3.6552 | 3.3628 |  | 3.6352 | 3.7559 | 3.4554
ST | 60 | ST-60 | 12 | 2.7874 | 2.8326 | 2.9887 | 3.1301 | 3.3025 | 3.7652 | 3.8611 |  | 3.5522 | 3.663 | 3.37 |  | 3.688 | 3.7755 | 3.4735
ST/Bone | 51 | ST/Bone-51 | 12 | 2.054 | 2.0702 | 2.3175 | 2.1298 | 2.3843 | 3.2456 | 3.1831 |  | 2.9285 | 3.0215 | 2.7798 |  | 3.0568 | 3.4053 | 3.1328
ST | 50 | ST-50 | 11 | 2.6223 | 2.7787 | 2.5673 | 3.3691 | 3.1128 | 3.8327 | 3.772 |  | 3.4702 | 3.9943 | 3.6747 |  | 3.7464 | 3.6505 | 3.3585
Bone | 39 | Bone-39 | 11 | 11.3743 | 7.2987 | 7.0849 | 10.4331 | 10.1274 | 11.4146 | 11.3051 |  | 10.4007 | 11.2838 | 10.3811 |  | 11.8578 | 10.755 | 9.8946
Bone | 37 | Bone-37 | 11 | 9.8673 | 6.8444 | 6.4017 | 8.5959 | 8.0399 | 11.2526 | 11.2357 |  | 10.3369 | 11.6535 | 10.7212 |  | 10.7543 | 11.0091 | 10.1284
Bone | 38 | Bone-38 | 11 | 7.4692 | 5.1928 | 5.1976 | 7.807 | 7.8142 | 10.1483 | 9.9694 |  | 9.1719 | 10.025 | 9.223 |  | 9.0457 | 9.5458 | 8.7821
ST/Bone | 44 | ST/Bone-44 | 11 | 3.2731 | 1.6319 | 1.578 | 2.5009 | 2.4184 | 3.4336 | 3.3967 |  | 3.125 | 3.5796 | 3.2932 |  | 3.3391 | 2.9981 | 2.7582
Lung | 48 | Lung-48 | 11 | 3.5979 | 2.4027 | 2.162 | 3.3596 | 3.023 | 4.2664 | 4.0722 |  | 3.7464 | 4.9524 | 4.5562 |  | 4.1625 | 4.4526 | 4.0964
Bone | 30 | Bone-30 | 10 | 9.7793 | 6.6081 | 5.4991 | 8.7249 | 7.2606 | 10.466 | 9.965 |  | 9.1678 | 10.9535 | 10.0772 |  | 10.2506 | 8.6071 | 7.9185
Bone | 33 | Bone-33 | 10 | 8.8333 | 7.424 | 6.3673 | 8.7862 | 7.5356 | 11.0248 | 11.3149 |  | 10.4097 | 12.0754 | 11.1093 |  | 10.8913 | 9.7409 | 8.9616
ST | 34 | ST-34 | 10 | 3.108 | 2.8879 | 2.3615 | 3.7043 | 3.0291 | 3.6719 | 3.4795 |  | 3.2012 | 3.8221 | 3.5163 |  | 3.4965 | 3.1282 | 2.8779
ST | 35 | ST-35 | 10 | 2.6667 | 2.1919 | 1.8967 | 2.7322 | 2.3642 | 3.1983 | 3.0531 |  | 2.8089 | 3.3122 | 3.0473 |  | 3.0288 | 2.8531 | 2.6249
Lung | 28 | Lung-28 | 10 | 3.2677 | 2.3448 | 1.9544 | 2.7055 | 2.2549 | 4.1225 | 3.9736 |  | 3.6557 | 4.2528 | 3.9126 |  | 3.6643 | 3.7978 | 3.494

## 5yrOLD-organwise
Scan Parameters
| HF - Head First Scan Position | 300 mAs | 160 mm Collimation
| FF - Feet First Scan Position | 120 Kvp | Body Filter
Voxel with Table but no FOAM (C density 1.2g/cc)
Organ | Measured Dose (cGy) (HF - run 1) | Measured Dose (cGy) (HF - run 2) | Measured Dose (cGy) (FF - run 3) | FLUKA HF | FLUKA FF
Bone | 5.5442 | 9.0031 | 10.9006 | 10.3503 | 11.007
Bone | 5.2472 | 10.1508 | 11.2379 | 9.8377 | 10.7347
Bone | 6.7322 | 10.8915 | 12.7299 | 11.4368 | 12.3908
Bone | 11.3743 | 10.4331 | 10.1274 | 11.4146 | 11.3051 | PLOTTING
Bone | 9.8673 | 8.5959 | 8.0399 | 11.2526 | 11.2357
Bone | 7.4692 | 7.807 | 7.8142 | 10.1483 | 9.9694 |  | Head First Scans
Bone | 9.7793 | 8.7249 | 7.2606 | 10.466 | 9.965 |  | Measured Dose |  | Fluka Dose |  | Fluka Dose (with CTDI correction)
Bone | 8.8333 | 8.7862 | 7.5356 | 11.0248 | 11.3149 |  | Mean | S.D | Mean | S.D | Mean | S.D
Avg | 8.1059 | 9.299 | 9.4558 | 10.7414 | 10.9903 | Bone | 9.299 | 1.0656 | 10.7414 | 0.61827 | 9.8821 | 0.56881
S.D | 2.2085 | 1.0656 | 2.0577 | 0.61827 | 0.7915 | Lung | 3.0007 | 0.54336 | 4.0008 | 0.22126 | 3.6807 | 0.20356
[6 empty] | Soft Tissue | 3.0338 | 0.38663 | 3.577 | 0.27544 | 3.2908 | 0.2534
Lung | 1.9675 | 2.8379 | 3.7385 | 3.9047 | 4.3269 | Soft Tissue/Bone | 2.3154 | 0.26243 | 3.3396 | 0.13296 | 3.0724 | 0.12232
Lung | 2.2839 | 2.6693 | 3.5663 | 3.8169 | 4.1921
Lung | 2.1842 | 2.6642 | 3.4546 | 4.0202 | 4.3576 |  | Feet First Scans
Lung | 1.964 | 2.3002 | 3.006 | 3.6274 | 4.1133 |  | Measured Dose |  | Fluka Dose |  | Fluka Dose (with CTDI correction)
Lung | 1.9812 | 3.6047 | 4.2639 | 3.9757 | 4.5181 |  | Mean | S.D | Mean | S.D | Mean | S.D
Lung | 2.5282 | 3.8641 | 4.4796 | 4.2725 | 4.461 | Bone | 9.4558 | 2.0577 | 10.9903 | 0.7915 | 10.1111 | 0.72818
Lung | 3.5979 | 3.3596 | 3.023 | 4.2664 | 4.0722 | Lung | 3.4733 | 0.71937 | 4.2519 | 0.19417 | 3.9117 | 0.17864
Lung | 3.2677 | 2.7055 | 2.2549 | 4.1225 | 3.9736 | Soft Tissue | 3.0264 | 0.37673 | 3.5872 | 0.27912 | 3.3002 | 0.25679
Avg | 2.4718 | 3.0007 | 3.4733 | 4.0008 | 4.2519 | Soft Tissue/Bone | 2.4013 | 0.024157 | 3.2899 | 0.15102 | 3.0267 | 0.13894
S.D | 0.62968 | 0.54336 | 0.71937 | 0.22126 | 0.19417
Soft. Tissue | 1.6984 | 2.6912 | 3.5707 | 3.1653 | 3.4719
Soft. Tissue | 2.7937 | 2.8833 | 2.9386 | 3.6814 | 3.7793
Soft. Tissue | 2.7937 | 2.7262 | 2.867 | 3.724 | 3.6933
Soft. Tissue | 2.7874 | 3.1301 | 3.3025 | 3.7652 | 3.8611
Soft. Tissue | 2.6223 | 3.3691 | 3.1128 | 3.8327 | 3.772
Soft. Tissue | 3.108 | 3.7043 | 3.0291 | 3.6719 | 3.4795
Soft. Tissue | 2.6667 | 2.7322 | 2.3642 | 3.1983 | 3.0531
Avg | 2.6386 | 3.0338 | 3.0264 | 3.577 | 3.5872
S.D | 0.44263 | 0.38663 | 0.37673 | 0.27544 | 0.27912
ST/Bone | 2.054 | 2.1298 | 2.3843 | 3.2456 | 3.1831
ST/Bone | 3.2731 | 2.5009 | 2.4184 | 3.4336 | 3.3967
Avg | 2.6635 | 2.3154 | 2.4013 | 3.3396 | 3.2899
S.D | 0.86201 | 0.26243 | 0.024157 | 0.13296 | 0.15102
Organ | Dose
Bone | 10.3503
Bone | 9.8377
Bone | 11.4368
Bone | 11.4146
Bone | 11.2526
Bone | 10.1483
Bone | 10.466
Bone | 11.0248
Lung | 3.9047
Lung | 3.8169
Lung | 4.0202
Lung | 3.6274
Lung | 3.9757
Lung | 4.2725
Lung | 4.2664
Lung | 4.1225
ST | 3.1653
ST | 3.6814
ST | 3.724
ST | 3.7652
ST | 3.8327
ST | 3.6719
ST | 3.1983
ST/Bone | 3.2456
ST/Bone | 3.4336

## HF-Axial-All Measurements
|  |  |  | Average Measured Dose (cGy) |  |  |  | 120 KVP
Organ | Hole | Organ-Hole | Slab | Run 1 (Head First) - Old Calibration | Run 2 (Head First) - New Calibration | Run 5 (Head First) - New calibration | Run 9 (HF-Axial-Large Body-NC) | 300 mAS
Lung | 81 | Lung-81 | 14 | 1.9675 | 2.8379 | 2.7954 | 2.773
Lung | 84 | Lung-84 | 14 | 2.2839 | 2.6693 | 2.5772 | 2.9765
Lung | 77 | Lung-77 | 14 | 2.1842 | 2.6642 | 2.6264 | 2.8082
Lung | 79 | Lung-79 | 14 | 1.964 | 2.3002 | 2.4694 | 2.52
ST | 85 | ST-85 | 14 | 1.6984 | 2.6912 | 2.7179 | 2.6148
Lung | 68 | Lung-68 | 13 | 1.9812 | 3.6047 | 3.5656 | 3.8541
Lung | 72 | Lung-72 | 13 | 2.5282 | 3.8641 | 3.5697 | 4.2004
Bone | 63 | Bone-63 | 13 | 5.5442 | 9.0031 | 8.7289 | 10.2607
Bone | 65 | Bone-65 | 13 | 5.2472 | 10.1508 | 9.1642 | 11.3909
Bone | 76 | Bone-76 | 13 | 6.7322 | 10.8915 | 10.4268 | 11.3336
ST | 58 | ST-58 | 12 | 2.7937 | 2.8833 | 2.8678 | 2.9251
ST | 59 | ST-59 | 12 | 2.7937 | 2.7262 | 2.5711 | 2.8883
ST | 60 | ST-60 | 12 | 2.7874 | 3.1301 | 3.18 | 3.0994
ST/Bone | 51 | ST/Bone-51 | 12 | 2.054 | 2.1298 | 2.1322 | 2.1402
ST | 50 | ST-50 | 11 | 2.6223 | 3.3691 | 3.3252 | 3.1678
Bone | 39 | Bone-39 | 11 | 11.3743 | 10.4331 | 10.6589 | 10.2683
Bone | 37 | Bone-37 | 11 | 9.8673 | 8.5959 | 8.1696 | 9.4975
Bone | 38 | Bone-38 | 11 | 7.4692 | 7.807 | 7.7669 | 7.7344
ST/Bone | 44 | ST/Bone-44 | 11 | 3.2731 | 2.5009 | 2.4944 | 2.67
Lung | 48 | Lung-48 | 11 | 3.5979 | 3.3596 | 3.2722 | 3.2955
Bone | 30 | Bone-30 | 10 | 9.7793 | 8.7249 | 8.8466 | 9.2105
Bone | 33 | Bone-33 | 10 | 8.8333 | 8.7862 | 8.9599 | 9.668
ST | 34 | ST-34 | 10 | 3.108 | 3.7043 | 3.6237 | 3.3482
ST | 35 | ST-35 | 10 | 2.6667 | 2.7322 | 2.6159 | 2.5322
Lung | 28 | Lung-28 | 10 | 3.2677 | 2.7055 | 2.6762 | 2.6691

## FF-Axial-All Measurements
120 KVP
Organ | Hole | Organ-Hole | Slab | Run 3 (Feet First) - New Calibration | Run 7 (Feet First) - New calibration | 300 mAS
Lung | 81 | Lung-81 | 14 | 3.7385 | 3.6669
Lung | 84 | Lung-84 | 14 | 3.5663 | 2.8367
Lung | 77 | Lung-77 | 14 | 3.4546 | 3.2057
Lung | 79 | Lung-79 | 14 | 3.006 | 2.6112
ST | 85 | ST-85 | 14 | 3.5707 | 3.0055
Lung | 68 | Lung-68 | 13 | 4.2639 | 2.906
Lung | 72 | Lung-72 | 13 | 4.4796 | 3.0584
Bone | 63 | Bone-63 | 13 | 10.9006 | 7.1541
Bone | 65 | Bone-65 | 13 | 11.2379 | 8.7431
Bone | 76 | Bone-76 | 13 | 12.7299 | 10.3352
ST | 58 | ST-58 | 12 | 2.9386 | 2.7096
ST | 59 | ST-59 | 12 | 2.867 | 2.658
ST | 60 | ST-60 | 12 | 3.3025 | 2.9171
ST/Bone | 51 | ST/Bone-51 | 12 | 2.3843 | 2.1542
ST | 50 | ST-50 | 11 | 3.1128 | 2.5389
Bone | 39 | Bone-39 | 11 | 10.1274 | 6.859
Bone | 37 | Bone-37 | 11 | 8.0399 | 6.0584
Bone | 38 | Bone-38 | 11 | 7.8142 | 5.0571
ST/Bone | 44 | ST/Bone-44 | 11 | 2.4184 | 1.4344
Lung | 48 | Lung-48 | 11 | 3.023 | 2.1079
Bone | 30 | Bone-30 | 10 | 7.2606 | 5.3296
Bone | 33 | Bone-33 | 10 | 7.5356 | 5.901
ST | 34 | ST-34 | 10 | 3.0291 | 2.4412
ST | 35 | ST-35 | 10 | 2.3642 | 1.9496
Lung | 28 | Lung-28 | 10 | 2.2549 | 2.0007

## HA Axial Runs - 5 yr Old
HA - High Accuracy Material Properties were used to generate the VOXEL |  |  |  |  | Particle Histories | 10^8
Low Accuracy Simulations (From sheet: 5 yr old - multiple runs)
|  |  |  | Measured Dose Centi Gy |  | Fluka Simulations with Table |  | Fluka Simulations without Table |  | Fluka Simlations
Organ | Hole | Organ-Hole | Slab | Run 2 (Head First) - New Calibration | Run 3 (Feet First) - New Calibration | Fluka Dose (Head First) | Fluka Dose (Feet First) | Fluka Dose (Head First) | Fluka Dose (Feet First) | Fluka Dose (Head First) | Fluka Dose (Feet First)
Lung | 81 | Lung-81 | 14 | 2.8379 | 3.7385 | 3.863 | 4.092 | 3.9698 | 4.418 | 3.9047 | 4.3269
Lung | 84 | Lung-84 | 14 | 2.6693 | 3.5663 | 3.6463 | 4.0318 | 3.63 | 4.4644 | 3.8169 | 4.1921
Lung | 77 | Lung-77 | 14 | 2.6642 | 3.4546 | 3.5175 | 4.2981 | 3.586 | 4.4943 | 4.0202 | 4.3576
Lung | 79 | Lung-79 | 14 | 2.3002 | 3.006 | 3.3955 | 4.1651 | 3.2712 | 3.9364 | 3.6274 | 4.1133
ST | 85 | ST-85 | 14 | 2.6912 | 3.5707 | 3.1096 | 3.5003 | 3.0633 | 3.6834 | 3.1653 | 3.4719
Lung | 68 | Lung-68 | 13 | 3.6047 | 4.2639 | 3.9994 | 4.1167 | 3.7801 | 4.4642 | 3.9757 | 4.5181
Lung | 72 | Lung-72 | 13 | 3.8641 | 4.4796 | 3.8744 | 4.0126 | 3.8833 | 4.3854 | 4.2725 | 4.461
Bone | 63 | Bone-63 | 13 | 9.0031 | 10.9006 | 10.6208 | 11.0266 | 10.9785 | 12.0384 | 10.3503 | 11.007
Bone | 65 | Bone-65 | 13 | 10.1508 | 11.2379 | 9.9445 | 10.4277 | 10.7009 | 11.3389 | 9.8377 | 10.7347
Bone | 76 | Bone-76 | 13 | 10.8915 | 12.7299 | 11.0168 | 11.4671 | 11.2259 | 11.9476 | 11.4368 | 12.3908
ST | 58 | ST-58 | 12 | 2.8833 | 2.9386 | 3.6536 | 3.5072 | 3.7199 | 3.6402 | 3.6814 | 3.7793
ST | 59 | ST-59 | 12 | 2.7262 | 2.867 | 3.6496 | 3.7096 | 3.7116 | 3.9968 | 3.724 | 3.6933
ST | 60 | ST-60 | 12 | 3.1301 | 3.3025 | 3.9096 | 3.9766 | 3.9841 | 4.2018 | 3.7652 | 3.8611
ST/Bone | 51 | ST/Bone-51 | 12 | 2.1298 | 2.3843 | 3.0648 | 3.1764 | 3.3297 | 3.3066 | 3.2456 | 3.1831
ST | 50 | ST-50 | 11 | 3.3691 | 3.1128 | 3.8585 | 3.7966 | 4.0374 | 3.9046 | 3.8327 | 3.772
Bone | 39 | Bone-39 | 11 | 10.4331 | 10.1274 | 11.2473 | 11.3069 | 11.7358 | 11.751 | 11.4146 | 11.3051
Bone | 37 | Bone-37 | 11 | 8.5959 | 8.0399 | 10.8344 | 11.1939 | 11.8969 | 11.9307 | 11.2526 | 11.2357
Bone | 38 | Bone-38 | 11 | 7.807 | 7.8142 | 9.8137 | 9.7194 | 10.4428 | 10.2782 | 10.1483 | 9.9694
ST/Bone | 44 | ST/Bone-44 | 11 | 2.5009 | 2.4184 | 3.5979 | 3.4014 | 3.6383 | 3.5469 | 3.4336 | 3.3967
Lung | 48 | Lung-48 | 11 | 3.3596 | 3.023 | 4.4261 | 3.8283 | 4.2425 | 4.54 | 4.2664 | 4.0722
Bone | 30 | Bone-30 | 10 | 8.7249 | 7.2606 | 10.4055 | 9.7487 | 10.8709 | 10.3209 | 10.466 | 9.965
Bone | 33 | Bone-33 | 10 | 8.7862 | 7.5356 | 11.0115 | 10.8646 | 11.7616 | 11.2508 | 11.0248 | 11.3149
ST | 34 | ST-34 | 10 | 3.7043 | 3.0291 | 3.5279 | 3.4786 | 3.834 | 3.59 | 3.6719 | 3.4795
ST | 35 | ST-35 | 10 | 2.7322 | 2.3642 | 3.2909 | 3.1878 | 3.3788 | 3.4248 | 3.1983 | 3.0531
Lung | 28 | Lung-28 | 10 | 2.7055 | 2.2549 | 3.7744 | 3.8621 | 3.8139 | 3.6395 | 4.1225 | 3.9736
[6 empty] | 5.8821 | 5.9958 | 6.0995 | 6.3398 | 5.9862 | 6.1451
Bone | 9.0031 | 10.9006 | 10.6208 | 11.0266 | 3.7385 | 3.7025
Bone | 10.1508 | 11.2379 | 9.9445 | 10.4277 | 3.5663 | 3.3182
Bone | 10.8915 | 12.7299 | 11.0168 | 11.4671 | 3.4546 | 3.474
Bone | 10.4331 | 10.1274 | 11.2473 | 11.3069 | 3.006 | 2.8592
Bone | 8.5959 | 8.0399 | 10.8344 | 11.1939 | 3.5707 | 3.4721
Bone | 7.807 | 7.8142 | 9.8137 | 9.7194 | 4.2639 | 3.4947
Bone | 8.7249 | 7.2606 | 10.4055 | 9.7487 | 4.4796 | 3.7828
Bone | 8.7862 | 7.5356 | 11.0115 | 10.8646 | 10.9006 | 9.3771
Lung | 2.8379 | 3.7385 | 3.863 | 4.092 | 11.2379 | 9.5581
Lung | 2.6693 | 3.5663 | 3.6463 | 4.0318 | 12.7299 | 11.0878
Lung | 2.6642 | 3.4546 | 3.5175 | 4.2981 | 2.9386 | 2.8488
Lung | 2.3002 | 3.006 | 3.3955 | 4.1651 | 2.867 | 2.6518
Lung | 3.6047 | 4.2639 | 3.9994 | 4.1167 | 3.3025 | 3.2234
Lung | 3.8641 | 4.4796 | 3.8744 | 4.0126 | 2.3843 | 2.2163
Lung | 3.3596 | 3.023 | 4.4261 | 3.8283 | 3.1128 | 3.0785
Lung | 2.7055 | 2.2549 | 3.7744 | 3.8621 | 10.1274 | 9.8045
ST | 2.6912 | 3.5707 | 3.1096 | 3.5003 | 8.0399 | 7.6087
ST | 2.8833 | 2.9386 | 3.6536 | 3.5072 | 7.8142 | 7.6029
ST | 2.7262 | 2.867 | 3.6496 | 3.7096 | 2.4184 | 2.1983
ST | 3.1301 | 3.3025 | 3.9096 | 3.9766 | 3.023 | 2.9473
ST | 3.3691 | 3.1128 | 3.8585 | 3.7966 | 7.2606 | 7.0368
ST | 3.7043 | 3.0291 | 3.5279 | 3.4786 | 7.5356 | 6.9838
ST | 2.7322 | 2.3642 | 3.2909 | 3.1878 | 3.0291 | 3.1313
ST/Bone | 2.1298 | 2.3843 | 3.0648 | 3.1764 | 2.3642 | 2.4301
ST/Bone | 2.5009 | 2.4184 | 3.5979 | 3.4014 | 2.2549 | 2.3084

## HA-Axial-organwise
Scan Parameters
| HF - Head First Scan Position | 300 mAs | 160 mm Collimation
| FF - Feet First Scan Position | 120 Kvp | Body Filter
Voxel with Table but no FOAM (C density 1.5g/cc)
Organ | Measured Dose (cGy) (HF - run 1) | Measured Dose (cGy) (HF - run 2) | Measured Dose (cGy) (FF - run 3) | FLUKA HF | FLUKA FF
Bone | 5.5442 | 9.0031 | 10.9006 | 10.6208 | 11.0266
Bone | 5.2472 | 10.1508 | 11.2379 | 9.9445 | 10.4277
Bone | 6.7322 | 10.8915 | 12.7299 | 11.0168 | 11.4671
Bone | 11.3743 | 10.4331 | 10.1274 | 11.2473 | 11.3069 | PLOTTING
Bone | 9.8673 | 8.5959 | 8.0399 | 10.8344 | 11.1939
Bone | 7.4692 | 7.807 | 7.8142 | 9.8137 | 9.7194 |  | Head First Scans
Bone | 9.7793 | 8.7249 | 7.2606 | 10.4055 | 9.7487 |  | Measured Dose |  | Fluka Dose |  | Fluka Dose (with CTDI correction) |  | Corr Fac Used | 0.86
Bone | 8.8333 | 8.7862 | 7.5356 | 11.0115 | 10.8646 |  | Mean | S.D | Mean | S.D | Mean | S.D
Avg | 8.1059 | 9.299 | 9.4558 | 10.6118 | 10.7194 | Bone | 9.299 | 1.0656 | 10.6118 | 0.52128 | 9.1261 | 0.47957
S.D | 2.2085 | 1.0656 | 2.0577 | 0.52128 | 0.68371 | Lung | 3.0007 | 0.54336 | 3.8121 | 0.31799 | 3.2784 | 0.29255
[6 empty] | Soft Tissue | 3.0338 | 0.38663 | 3.5714 | 0.28968 | 3.0714 | 0.26651
Lung | 1.9675 | 2.8379 | 3.7385 | 3.863 | 4.092 | Soft Tissue/Bone | 2.3154 | 0.26243 | 3.3313 | 0.37698 | 2.865 | 0.34682
Lung | 2.2839 | 2.6693 | 3.5663 | 3.6463 | 4.0318
Lung | 2.1842 | 2.6642 | 3.4546 | 3.5175 | 4.2981 |  | Feet First Scans
Lung | 1.964 | 2.3002 | 3.006 | 3.3955 | 4.1651 |  | Measured Dose |  | Fluka Dose |  | Fluka Dose (with CTDI correction)
Lung | 1.9812 | 3.6047 | 4.2639 | 3.9994 | 4.1167 |  | Mean | S.D | Mean | S.D | Mean | S.D
Lung | 2.5282 | 3.8641 | 4.4796 | 3.8744 | 4.0126 | Bone | 9.4558 | 2.0577 | 10.7194 | 0.68371 | 9.2187 | 0.62901
Lung | 3.5979 | 3.3596 | 3.023 | 4.4261 | 3.8283 | Lung | 3.4733 | 0.71937 | 4.0508 | 0.15454 | 3.4837 | 0.14218
Lung | 3.2677 | 2.7055 | 2.2549 | 3.7744 | 3.8621 | Soft Tissue | 3.0264 | 0.37673 | 3.5938 | 0.25672 | 3.0907 | 0.23618
Avg | 2.4718 | 3.0007 | 3.4733 | 3.8121 | 4.0508 | Soft Tissue/Bone | 2.4013 | 0.024157 | 3.2889 | 0.15914 | 2.8284 | 0.14641
S.D | 0.62968 | 0.54336 | 0.71937 | 0.31799 | 0.15454
Soft. Tissue | 1.6984 | 2.6912 | 3.5707 | 3.1096 | 3.5003
Soft. Tissue | 2.7937 | 2.8833 | 2.9386 | 3.6536 | 3.5072
Soft. Tissue | 2.7937 | 2.7262 | 2.867 | 3.6496 | 3.7096
Soft. Tissue | 2.7874 | 3.1301 | 3.3025 | 3.9096 | 3.9766
Soft. Tissue | 2.6223 | 3.3691 | 3.1128 | 3.8585 | 3.7966
Soft. Tissue | 3.108 | 3.7043 | 3.0291 | 3.5279 | 3.4786
Soft. Tissue | 2.6667 | 2.7322 | 2.3642 | 3.2909 | 3.1878
Avg | 2.6386 | 3.0338 | 3.0264 | 3.5714 | 3.5938
S.D | 0.44263 | 0.38663 | 0.37673 | 0.28968 | 0.25672
ST/Bone | 2.054 | 2.1298 | 2.3843 | 3.0648 | 3.1764
ST/Bone | 3.2731 | 2.5009 | 2.4184 | 3.5979 | 3.4014
Avg | 2.6635 | 2.3154 | 2.4013 | 3.3313 | 3.2889
S.D | 0.86201 | 0.26243 | 0.024157 | 0.37698 | 0.15914

## HA-Helical Runs
| Head First Simulations | [5 empty] | Feet First Simulations
Organs | Measured | 0 deg | 90 deg | 180 deg | 270 deg | Average | Measured | 0 deg | 90 deg | 180 deg | 270 deg | Average
Lung | 1.37 | 1.2328 | 1.4178 | 2.0001 | 1.5979 | 1.5622 | 1.35 | 1.4853 | 1.9268 | 1.5983 | 1.9097 | 1.73
Lung | 1.23 | 1.2525 | 1.1249 | 1.5937 | 1.3847 | 1.339 | 0.99 | 1.5349 | 1.8616 | 1.4203 | 1.5759 | 1.5982
Lung | 0.94 | 1.6109 | 1.7434 | 1.6672 | 1.6581 | 1.6699 | 1.24 | 1.2801 | 1.2413 | 1.8732 | 1.4833 | 1.4695
Lung | 0.96 | 1.6459 | 1.5243 | 1.2108 | 1.1918 | 1.3932 | 1.06 | 1.6636 | 1.2074 | 1.4249 | 1.335 | 1.4078
ST | 1.16 | 1.2162 | 1.2702 | 1.3543 | 1.4306 | 1.3178 | 1.11 | 1.324 | 1.3078 | 1.3625 | 1.361 | 1.3388
Lung | 1.26 | 1.3336 | 1.9007 | 1.7806 | 1.722 | 1.6842 | 1.31 | 1.7542 | 1.4143 | 1.2471 | 1.1036 | 1.3798
Lung | 1.14 | 1.7147 | 1.1705 | 1.5048 | 1.1134 | 1.3759 | 1.12 | 1.1609 | 1.6015 | 1.8166 | 1.8298 | 1.6022
Bone | 3.36 | 3.9181 | 4.8765 | 3.9527 | 4.6138 | 4.3403 | 3.08 | 4.784 | 4.1651 | 3.3644 | 3.7567 | 4.0176
Bone | 3.88 | 4.5578 | 3.1291 | 2.9291 | 2.6229 | 3.3097 | 3.26 | 2.7158 | 4.5003 | 4.7465 | 4.8002 | 4.1907
Bone | 3.01 | 2.9019 | 3.2407 | 5.0871 | 4.2129 | 3.8606 | 3.67 | 3.7136 | 2.7691 | 4.7483 | 3.632 | 3.7157
ST | 1.1 | 1.1676 | 0.76507 | 1.4456 | 1.0345 | 1.1032 | 0.83 | 1.7004 | 1.2444 | 0.77232 | 0.90058 | 1.1544
ST | 0.76 | 1.6658 | 1.0583 | 0.84656 | 0.88232 | 1.1132 | 0.88 | 1.3955 | 0.85511 | 1.2418 | 0.93925 | 1.1079
ST | 1.1 | 1.458 | 1.1196 | 1.3045 | 1.0432 | 1.2313 | 0.93 | 1.6029 | 1.1953 | 1.1834 | 1.1063 | 1.272
ST/Bone | 0.87 | 1.1335 | 1.1746 | 1.1896 | 1.143 | 1.1602 | 0.84 | 1.2563 | 1.2008 | 1.2413 | 1.2558 | 1.2385
ST | 1.1 | 1.6665 | 1.3825 | 1.0815 | 1.1798 | 1.3276 | 1.05 | 1.6009 | 1.5633 | 1.1248 | 1.3044 | 1.3983
Bone | 3.13 | 5.1329 | 4.1081 | 2.6717 | 3.2754 | 3.797 | 2.63 | 5.0599 | 4.6691 | 2.9682 | 3.9083 | 4.1514
Bone | 2.18 | 3.5997 | 4.9585 | 4.3869 | 4.8361 | 4.4453 | 2.39 | 4.4167 | 2.7593 | 3.7625 | 2.9435 | 3.4705
Bone | 1.91 | 3.3494 | 3.5615 | 3.9817 | 3.7313 | 3.656 | 2.02 | 3.4172 | 3.4351 | 3.7082 | 3.5545 | 3.5287
ST/Bone | 0.56 | 1.265 | 1.448 | 1.1588 | 1.4357 | 1.3269 | 0.6 | 1.4562 | 0.97285 | 0.98829 | 0.85259 | 1.0675
Lung | 0.93 | 1.6103 | 1.7807 | 1.2843 | 1.4609 | 1.5341 | 0.88 | 1.5556 | 1.5705 | 1.2554 | 1.3695 | 1.4378
Bone | 1.9 | 4.7502 | 3.6041 | 2.2145 | 2.5291 | 3.2745 | 2.37 | 2.1641 | 3.5918 | 4.6629 | 4.3438 | 3.6907
Bone | 3.21 | 2.3146 | 4.4458 | 5.0569 | 5.0931 | 4.2276 | 2.66 | 5.158 | 4.5891 | 2.3203 | 3.1998 | 3.8168
ST | 1.11 | 1.4144 | 1.522 | 1.2773 | 1.386 | 1.3999 | 1.04 | 1.1663 | 1.4503 | 1.4137 | 1.5447 | 1.3938
ST | 0.8 | 1.1937 | 1.2295 | 1.1302 | 1.1765 | 1.1825 | 0.84 | 1.1507 | 1.2782 | 1.246 | 1.1846 | 1.2149
Lung | 0.87 | 1.5412 | 1.2234 | 1.251 | 1.361 | 1.3442 | 0.81 | 1.2623 | 1.1633 | 1.4483 | 1.3985 | 1.3181
S.D
|  | 0 deg | 90 deg | 180 deg | 270 deg | Average
|  | 0.079029 | 0.079403 | 0.10746 | 0.088354 | 0.044654
|  | 0.088373 | 0.07546 | 0.094723 | 0.091724 | 0.043939
|  | 0.10183 | 0.095183 | 0.097009 | 0.10455 | 0.049855
|  | 0.096355 | 0.088541 | 0.088921 | 0.074297 | 0.043697
|  | 0.036045 | 0.032581 | 0.041797 | 0.038918 | 0.018746
|  | 0.086246 | 0.10928 | 0.11475 | 0.098799 | 0.051424
|  | 0.091276 | 0.078541 | 0.091502 | 0.068598 | 0.041517
|  | 0.056711 | 0.069534 | 0.065666 | 0.064515 | 0.032138
|  | 0.071252 | 0.060717 | 0.050879 | 0.061189 | 0.030717
|  | 0.058447 | 0.048947 | 0.068701 | 0.069012 | 0.030918
|  | 0.029619 | 0.026007 | 0.037468 | 0.038382 | 0.016641
|  | 0.040733 | 0.032857 | 0.029762 | 0.030225 | 0.016841
|  | 0.033842 | 0.03644 | 0.033299 | 0.034573 | 0.017279
|  | 0.026565 | 0.037566 | 0.035076 | 0.028289 | 0.0161
|  | 0.041279 | 0.034466 | 0.033935 | 0.037188 | 0.018416
|  | 0.074267 | 0.062238 | 0.056874 | 0.067706 | 0.032795
|  | 0.060266 | 0.066648 | 0.066409 | 0.073403 | 0.033422
|  | 0.05979 | 0.056371 | 0.071778 | 0.06335 | 0.031542
|  | 0.034989 | 0.035627 | 0.033012 | 0.03976 | 0.017966
|  | 0.095549 | 0.10264 | 0.080708 | 0.077341 | 0.044833
|  | 0.067373 | 0.068207 | 0.048544 | 0.0569 | 0.030399
|  | 0.05373 | 0.076735 | 0.068733 | 0.068869 | 0.033766
|  | 0.039499 | 0.037962 | 0.035459 | 0.038538 | 0.018947
|  | 0.03322 | 0.037389 | 0.029465 | 0.03119 | 0.016474
|  | 0.093333 | 0.081549 | 0.098239 | 0.095606 | 0.046201

## HA-Helical Organ Wise
Organs | HF Measured | FF Measured | HF FLUKA - Average | FF FLUKA - Average
Bone | 3.36 | 3.08 | 4.3403 | 4.0176
Bone | 3.88 | 3.26 | 3.3097 | 4.1907
Bone | 3.01 | 3.67 | 3.8606 | 3.7157
Bone | 3.13 | 2.63 | 3.797 | 4.1514
Bone | 2.18 | 2.39 | 4.4453 | 3.4705
Bone | 1.91 | 2.02 | 3.656 | 3.5287
Bone | 1.9 | 2.37 | 3.2745 | 3.6907
Bone | 3.21 | 2.66 | 4.2276 | 3.8168
Lung | 1.37 | 1.35 | 1.5622 | 1.73
Lung | 1.23 | 0.99 | 1.339 | 1.5982
Lung | 0.94 | 1.24 | 1.6699 | 1.4695
Lung | 0.96 | 1.06 | 1.3932 | 1.4078
Lung | 1.26 | 1.31 | 1.6842 | 1.3798
Lung | 1.14 | 1.12 | 1.3759 | 1.6022
Lung | 0.93 | 0.88 | 1.5341 | 1.4378
Lung | 0.87 | 0.81 | 1.3442 | 1.3181
ST | 1.16 | 1.11 | 1.3178 | 1.3388
ST | 1.1 | 0.83 | 1.1032 | 1.1544
ST | 0.76 | 0.88 | 1.1132 | 1.1079
ST | 1.1 | 0.93 | 1.2313 | 1.272
ST | 1.1 | 1.05 | 1.3276 | 1.3983
ST | 1.11 | 1.04 | 1.3999 | 1.3938
ST | 0.8 | 0.84 | 1.1825 | 1.2149
ST/Bone | 0.87 | 0.84 | 1.1602 | 1.2385
ST/Bone | 0.56 | 0.6 | 1.3269 | 1.0675
| 300 mAs | 160 mm Collimation
| 120 Kvp | Body Filter
Voxel with Table but no FOAM (C density 1.5g/cc)
Organ | Measured Dose (cGy) (HF - Helical) | Measured Dose (cGy) (FF - Helical) | FLUKA HF | FLUKA FF
Bone | 3.36 | 3.08 | 4.3403 | 4.0176
Bone | 3.88 | 3.26 | 3.3097 | 4.1907
Bone | 3.01 | 3.67 | 3.8606 | 3.7157
Bone | 3.13 | 2.63 | 3.797 | 4.1514 | PLOTTING
Bone | 2.18 | 2.39 | 4.4453 | 3.4705
Bone | 1.91 | 2.02 | 3.656 | 3.5287 |  | Head First Scans
Bone | 1.9 | 2.37 | 3.2745 | 3.6907 |  | Measured Dose |  | Fluka Dose |  | Fluka Dose (with CTDI correction) |  | Corr Fac Used | 0.86
Bone | 3.21 | 2.66 | 4.2276 | 3.8168 |  | Mean | S.D | Mean | S.D | Mean | S.D
Avg | 2.8225 | 2.76 | 3.8639 | 3.8228 | Bone | 2.8225 | 0.73519 | 3.8639 | 0.44714 | 3.3229 | 0.41137
S.D | 0.73519 | 0.54037 | 0.44714 | 0.27276 | Lung | 1.0875 | 0.18622 | 1.4878 | 0.14327 | 1.2795 | 0.13181
[5 empty] | Soft Tissue | 1.0186 | 0.16477 | 1.2394 | 0.11363 | 1.0659 | 0.10454
Lung | 1.37 | 1.35 | 1.5622 | 1.73 | Soft Tissue/Bone | 0.715 | 0.2192 | 1.2435 | 0.11787 | 1.0694 | 0.10844
Lung | 1.23 | 0.99 | 1.339 | 1.5982
Lung | 0.94 | 1.24 | 1.6699 | 1.4695 |  | Feet First Scans
Lung | 0.96 | 1.06 | 1.3932 | 1.4078 |  | Measured Dose |  | Fluka Dose |  | Fluka Dose (with CTDI correction)
Lung | 1.26 | 1.31 | 1.6842 | 1.3798 |  | Mean | S.D | Mean | S.D | Mean | S.D
Lung | 1.14 | 1.12 | 1.3759 | 1.6022 | Bone | 2.76 | 0.54037 | 3.8228 | 0.27276 | 3.2876 | 0.25094
Lung | 0.93 | 0.88 | 1.5341 | 1.4378 | Lung | 1.095 | 0.19734 | 1.4929 | 0.13807 | 1.2839 | 0.12703
Lung | 0.87 | 0.81 | 1.3442 | 1.3181 | Soft Tissue | 0.95429 | 0.11208 | 1.2686 | 0.11487 | 1.091 | 0.10568
Avg | 1.0875 | 1.095 | 1.4878 | 1.4929 | Soft Tissue/Bone | 0.72 | 0.16971 | 1.153 | 0.12096 | 0.9916 | 0.11128
S.D | 0.18622 | 0.19734 | 0.14327 | 0.13807
Soft. Tissue | 1.16 | 1.11 | 1.3178 | 1.3388
Soft. Tissue | 1.1 | 0.83 | 1.1032 | 1.1544
Soft. Tissue | 0.76 | 0.88 | 1.1132 | 1.1079
Soft. Tissue | 1.1 | 0.93 | 1.2313 | 1.272
Soft. Tissue | 1.1 | 1.05 | 1.3276 | 1.3983
Soft. Tissue | 1.11 | 1.04 | 1.3999 | 1.3938
Soft. Tissue | 0.8 | 0.84 | 1.1825 | 1.2149
Avg | 1.0186 | 0.95429 | 1.2394 | 1.2686
S.D | 0.16477 | 0.11208 | 0.11363 | 0.11487
ST/Bone | 0.87 | 0.84 | 1.1602 | 1.2385
ST/Bone | 0.56 | 0.6 | 1.3269 | 1.0675
Avg | 0.715 | 0.72 | 1.2435 | 1.153
S.D | 0.2192 | 0.16971 | 0.11787 | 0.12096

## AdultFemale-Results
Phantom Location |  |  |  | Slice # |  | Measurement |  | FLUKA Results |  |  |  |  | Fluka Corrected |  | Correction
Organ | Tissue | Tissue Index | Hole |  | Tissue/Hole | Actual Dose (cGy) | S.D | 0 deg | 90 deg | 180 deg | 270 deg | Average | Dose | S.D | 0.9
Ovaries | Soft Tissue | ST | 212 | 31 | ST -212 | 0.29829 | 0.0098432 | 0.48876 | 0.47793 | 0.46292 | 0.51982 | 0.48736 | 0.43862 | 0.012564
Ovaries | Soft Tissue | ST | 213 | 31 | ST -213 | 0.34999 | 0.03473 | 0.44826 | 0.44669 | 0.45351 | 0.45204 | 0.45013 | 0.40511 | 0.012326
Bladder | Soft Tissue | ST | 215 | 31 | ST -215 | 0.47309 | 0.030479 | 0.69928 | 0.64906 | 0.65807 | 0.66992 | 0.66908 | 0.60218 | 0.014952
Bladder | Soft Tissue | ST | 225 | 32 | ST -225 | 0.40755 | 0.016532 | 0.5785 | 0.5218 | 0.57213 | 0.58445 | 0.56422 | 0.5078 | 0.013326
Bladder | Soft Tissue | ST | 224 | 32 | ST -224 | 0.41622 | 0.028134 | 1.0332 | 0.99936 | 1.0648 | 0.93649 | 1.0085 | 0.90761 | 0.019339
Bladder | Soft Tissue | ST | 234 | 33 | ST -234 | 0.37374 | 0.026133 | 0.53729 | 0.52824 | 0.45642 | 0.52042 | 0.51059 | 0.45953 | 0.013573
Kidney | Soft Tissue | ST | 179 | 26 | ST -179 | 0.49689 | 0.022488 | 0.66946 | 0.56779 | 0.62361 | 0.54772 | 0.60214 | 0.54193 | 0.013963
Liver | Soft Tissue | ST | 168 | 25 | ST -168 | 0.35838 | 0.015217 | 0.62902 | 0.56961 | 0.60622 | 0.56897 | 0.59346 | 0.53411 | 0.014602
Kidney | Soft Tissue | ST | 176 | 25 | ST -176 | 0.38866 | 0.014781 | 0.50422 | 0.52034 | 0.51432 | 0.58682 | 0.53142 | 0.47828 | 0.012821
stomach | Soft Tissue | ST | 159 | 24 | ST -159 | 0.33291 | 0.011716 | 0.48878 | 0.48225 | 0.46527 | 0.4807 | 0.47925 | 0.43133 | 0.01275
Liver | Soft Tissue | ST | 153 | 24 | ST -153 | 0.35769 | 0.018786 | 0.53696 | 0.54039 | 0.49425 | 0.50355 | 0.51879 | 0.46691 | 0.013002
Pancreas | Soft Tissue | ST | 162 | 24 | ST -162 | 0.34485 | 0.018731 | 0.44018 | 0.50605 | 0.41587 | 0.47624 | 0.45958 | 0.41363 | 0.011387
Kidney | Soft Tissue | ST | 164 | 24 | ST -164 | 0.38988 | 0.010417 | 0.48748 | 0.5451 | 0.49599 | 0.49395 | 0.50563 | 0.45507 | 0.012891
Kidney | Soft Tissue | ST | 150 | 23 | ST -150 | 0.31894 | 0.010665 | 0.45559 | 0.45841 | 0.4376 | 0.47139 | 0.45575 | 0.41017 | 0.012493
Liver | Soft Tissue | ST | 130 | 22 | ST -130 | 0.2932 | 0.027492 | 0.41891 | 0.43519 | 0.38162 | 0.41325 | 0.41224 | 0.37102 | 0.011801
Spleen | Soft Tissue | ST | 126 | 21 | ST -126 | 0.45187 | 0.018383 | 0.37945 | 0.40639 | 0.46258 | 0.42392 | 0.41809 | 0.37628 | 0.011552
Lung | Lung | Lung | 98 | 20 | Lung -98 | 0.53446 | 0.010095 | 0.45287 | 0.56165 | 0.49754 | 0.4639 | 0.49399 | 0.44459 | 0.028146
LIver | Soft Tissue | ST | 110 | 20 | ST -110 | 0.39084 | 0.0050082 | 0.41675 | 0.44479 | 0.44512 | 0.4339 | 0.43514 | 0.39163 | 0.011047
Lung | Lung | Lung | 85 | 18 | Lung -85 | 0.44968 | 0.0053958 | 0.51208 | 0.5297 | 0.36319 | 0.44723 | 0.46305 | 0.41674 | 0.030604
Lung | Lung | Lung | 61 | 16 | Lung -61 | 0.40191 | 0.019554 | 0.49858 | 0.47811 | 0.43484 | 0.41755 | 0.45727 | 0.41154 | 0.028424
Esophagus | Soft Tissue | ST | 60 | 15 | ST -60 | 0.35205 | 0.0092012 | 0.3982 | 0.43902 | 0.46733 | 0.38151 | 0.42152 | 0.37937 | 0.011205
Lung | Lung | Lung | 44 | 14 | Lung -44 | 0.42125 | 0.012525 | 0.50575 | 0.44237 | 0.52539 | 0.53065 | 0.50104 | 0.45094 | 0.031987
Thyroid | Soft Tissue | ST | 16 | 10 | ST -16 | 0.4332 | 0.075519 | 0.487 | 0.53773 | 0.523 | 0.86876 | 0.60413 | 0.54371 | 0.056367
Breasts | Breast | Breast | 243 | - | Breast -243 | 0.31218 | 0.040062 | 0.27235 | 0.31405 | 0.33945 | 0.36486 | 0.32268 | 0.29041 | 0.009925
Breasts | Breast | Breast | 247 | - | Breast -247 | 0.34884 | 0.010957 | 0.20266 | 0.25692 | 0.36237 | 0.42677 | 0.31218 | 0.28096 | 0.089038
FLUKA Errors
[8 empty] | 0deg | 90deg | 180 deg | 270deg | Error
[8 empty] | 0.026754 | 0.028239 | 0.024047 | 0.032044 | 0.013961
[8 empty] | 0.029331 | 0.027433 | 0.0262 | 0.026488 | 0.013695
[8 empty] | 0.032294 | 0.034828 | 0.036526 | 0.028736 | 0.016613
[8 empty] | 0.029356 | 0.030741 | 0.027885 | 0.030387 | 0.014807
[8 empty] | 0.045704 | 0.040589 | 0.044104 | 0.041304 | 0.021488
[8 empty] | 0.031871 | 0.029625 | 0.026911 | 0.031962 | 0.015081
[8 empty] | 0.03272 | 0.031356 | 0.029408 | 0.030541 | 0.015515
[8 empty] | 0.031385 | 0.033325 | 0.034028 | 0.030957 | 0.016225
[8 empty] | 0.029726 | 0.027381 | 0.027493 | 0.029285 | 0.014245
[8 empty] | 0.02883 | 0.02785 | 0.029508 | 0.027082 | 0.014166
[8 empty] | 0.028268 | 0.02975 | 0.030964 | 0.026389 | 0.014447
[8 empty] | 0.025771 | 0.025558 | 0.024073 | 0.025779 | 0.012653
[8 empty] | 0.027863 | 0.03298 | 0.028213 | 0.024946 | 0.014323
[8 empty] | 0.02762 | 0.026278 | 0.029075 | 0.028008 | 0.013882
[8 empty] | 0.024349 | 0.028714 | 0.023928 | 0.027584 | 0.013112
[8 empty] | 0.024071 | 0.026117 | 0.027794 | 0.024535 | 0.012835
[8 empty] | 0.065956 | 0.063837 | 0.0659 | 0.053664 | 0.031273
[8 empty] | 0.024425 | 0.026251 | 0.021405 | 0.025817 | 0.012274
[8 empty] | 0.071988 | 0.076489 | 0.056909 | 0.065031 | 0.034004
[8 empty] | 0.066331 | 0.060529 | 0.063956 | 0.061681 | 0.031582
[8 empty] | 0.023497 | 0.028444 | 0.022776 | 0.024501 | 0.01245
[8 empty] | 0.07429 | 0.057393 | 0.074974 | 0.076005 | 0.035541
[8 empty] | 0.029867 | 0.031302 | 0.07947 | 0.23361 | 0.06263
[8 empty] | 0.021962 | 0.021964 | 0.021458 | 0.022817 | 0.011028
[8 empty] | 0.11026 | 0.10777 | 0.25012 | 0.26508 | 0.098931

## AdultFemale-Organwise
| Measurement |  |  | FLUKA
Tissue/Hole | Actual Dose (cGy) | S.D |  | Average | Error | Dose
Breast -243 | 0.31218 | 0.040062 |  | 0.32268 | 0.011028 |  | Helical
Breast -247 | 0.34884 | 0.010957 |  | 0.31218 | 0.098931 |  | Breasts | Lungs | Soft Tissue
Mean | 0.33051 | 0.020767 | Mean | 0.31743 | 0.049772 | Measured | 0.33051 | 0.45183 | 0.38043
SD | 0.018328 |  | SD | 0.0052503 |  | FLUKA | 0.31743 | 0.47884 | 0.533
FLUKA (CTDI Correction)
Lung -44 | 0.42125 | 0.012525 |  | 0.50104 | 0.035541
Lung -61 | 0.40191 | 0.019554 |  | 0.45727 | 0.031582 | S.D
Lung -85 | 0.44968 | 0.0053958 |  | 0.46305 | 0.034004 |  | Helical
Lung -98 | 0.53446 | 0.010095 |  | 0.49399 | 0.031273 |  | Breasts | Lungs | Soft Tissue
Mean | 0.45183 | 0.0064724 | Mean | 0.47884 | 0.016573 | Measured | 0.020767 | 0.0064724 | 0.050556
SD | 0.029238 |  | SD | 0.010943 |  | FLUKA | 0.049772 | 0.016573 | 0.049759
FLUKA (CTDI Correction)
ST -110 | 0.39084 | 0.0050082 |  | 0.43514 | 0.012274
ST -126 | 0.45187 | 0.018383 |  | 0.41809 | 0.012835 | 95% CI
ST -130 | 0.2932 | 0.027492 |  | 0.41224 | 0.013112 |  | Helical
ST -150 | 0.31894 | 0.010665 |  | 0.45575 | 0.013882 |  | Breasts | Lungs | Soft Tissue
ST -153 | 0.35769 | 0.018786 |  | 0.51879 | 0.014447 | Measured | 0.035922 | 0.057307 | 0.02525
ST -159 | 0.33291 | 0.011716 |  | 0.47925 | 0.014166 | FLUKA | 0.010291 | 0.021448 | 0.061205
ST -16 | 0.4332 | 0.075519 |  | 0.60413 | 0.06263 | FLUKA (CTDI Correction)
ST -162 | 0.34485 | 0.018731 |  | 0.45958 | 0.012653
ST -164 | 0.38988 | 0.010417 |  | 0.50563 | 0.014323
ST -168 | 0.35838 | 0.015217 |  | 0.59346 | 0.016225 | S.D Variability
ST -176 | 0.38866 | 0.014781 |  | 0.53142 | 0.014245 |  | Helical
ST -179 | 0.49689 | 0.022488 |  | 0.60214 | 0.015515 |  | Breasts | Lungs | Soft Tissue
ST -212 | 0.29829 | 0.0098432 |  | 0.48736 | 0.013961 | Measured | 0.018328 | 0.029238 | 0.012883
ST -213 | 0.34999 | 0.03473 |  | 0.45013 | 0.013695 | FLUKA | 0.0052503 | 0.010943 | 0.031227
ST -215 | 0.47309 | 0.030479 |  | 0.66908 | 0.016613 | FLUKA (CTDI Correction)
ST -224 | 0.41622 | 0.028134 |  | 1.0085 | 0.021488
ST -225 | 0.40755 | 0.016532 |  | 0.56422 | 0.014807
ST -234 | 0.37374 | 0.026133 |  | 0.51059 | 0.015081
ST -60 | 0.35205 | 0.0092012 |  | 0.42152 | 0.01245
Mean | 0.38043 | 0.050556 | Mean | 0.533 | 0.049759
SD | 0.012883 |  | SD | 0.031227
Tissue/Hole | Actual Dose (cGy) | S.D
Breast -243 | 0.31218 | 0.040062
Breast -247 | 0.34884 | 0.010957
Lung -44 | 0.42125 | 0.012525
Lung -61 | 0.40191 | 0.019554
Lung -85 | 0.44968 | 0.0053958
Lung -98 | 0.53446 | 0.010095
ST -110 | 0.39084 | 0.0050082
ST -126 | 0.45187 | 0.018383
ST -130 | 0.2932 | 0.027492
ST -150 | 0.31894 | 0.010665
ST -153 | 0.35769 | 0.018786
ST -159 | 0.33291 | 0.011716
ST -16 | 0.4332 | 0.075519
ST -162 | 0.34485 | 0.018731
ST -164 | 0.38988 | 0.010417
ST -168 | 0.35838 | 0.015217
ST -176 | 0.38866 | 0.014781
ST -179 | 0.49689 | 0.022488
ST -212 | 0.29829 | 0.0098432
ST -213 | 0.34999 | 0.03473
ST -215 | 0.47309 | 0.030479
ST -224 | 0.41622 | 0.028134
ST -225 | 0.40755 | 0.016532
ST -234 | 0.37374 | 0.026133
ST -60 | 0.35205 | 0.0092012